EANM Practice Guideline/SNMMI Procedure Standard for Dopaminergic Imaging in Parkinsonian Syndromes 1.0

Silvia Morbelli1,2, Giuseppe Esposito3, Javier Arbizu4, Henryk Barthel5, Ronald Boellaard6, Nico I. Bohnen7, David J Brooks8,9, Jacques Darcourt10, John C. Dickson11, David Douglas12, Alexander Drzezga13, Jacob Dubroff14, Ozgul Ekmekcioglu15, Valentina Garibotto16,17, Peter Herscovitch18, Phillip Kuo19, Adriaan Lammertsma6, Sabina Pappata20, Iván Peñuelas4, John Seibyl21, Franck Semah22, Livia Tossici-Bolt23, Elsmarieke Van de Giessen6, Koen Van Laere24, Andrea Varrone25, Michele Wanner26, George Zubal27, Ian Law28.

1. IRCCS Ospedale Policlinico San Martino, Genova, Italy
2. Nuclear Medicine Unit, Department of Health Sciences, University of Genoa, Italy
3. Nuclear Medicine, Department of Radiology, Medstar Georgetown University Hospital, 3800 Reservoir Road NW CCC Bldg., Washington, DC, 20007, USA.
4. Department of Nuclear Medicine. University of Navarra. Clinica Universidad de Navarra. IdiSNA. Pamplona, Spain

Department of Nuclear Medicine, Leipzig University, Leipzig, Germany

Department of Radiology and Nuclear Medicine, Amsterdam UMC, Vrije Universiteit Amsterdam, Cancer Center Amsterdam NeuroScience, De Boelelaan 1117, Amsterdam, the Netherlands

Departments Radiology & Neurology, University of Michigan & Ann Arbor Veterans Medical Center, Ann Arbor, MI, USA

1. Dpt of Nuclear Medicine Aarhus University DK;
2. Institute of Neuroscience, University of Newcastle Upon Tyne, UK

Nuclear Medicine. Department of imaging. Centre Antoine Lacassagne, Nice, Université Cote d’Azur, France

 Institute of Nuclear Medicine, University College London Hospitals NHS Foundation Trust, London

1. Department of Nuclear Medicine and Neuroradiology, Stanford University, CA, USA
2. Klinik und Poliklinik für Nuklearmedizin, Universität zu Köln, Köln, Germany
3. Division of Nuclear Medicine and Clinical Molecular Imaging in the Department of Radiology, Perelman School of Medicine at the University of Pennsylvania, Philadelphia, PA, USA
4. Department of Nuclear Medicine, Bagcilar Education and Research Hospital, Istanbul, Turkey
5. NIMTLab, Faculty of Medicine, Geneva University, Geneva, Switzerland
6. Division of Nuclear Medicine and Molecular Imaging, Geneva University Hospitals, Geneva, Switzerland
7. Positron Emission Tomography Department, NIH/CC, Bethesda, MD, USA.
8. Departments of Medical Imaging, Medicine and Biomedical Engineering, University of Arizona, USA.
9. Institute of Biostructure and Bioimaging, National Council of Research, Naples, Italy
10. Institute for Neurodegenerative Disorders, New Haven, USA
11. Nuclear Medicine Department, University Hospital, Lille, France

Medical Physics, University Hospital Southampton NHS Trust, Southampton, United Kingdom

Nuclear Medicine and PET-MR Unit, Department of Imaging and Pathology, University Hospital and KU Leuven, Leuven, Belgium

Department of Clinical Neuroscience, Centre for Psychiatry Research, Karolinska Institutet and Stockholm County Council, Stockholm, Sweden

1. University of Washington, Seattle, Washington
2. Nuclear Medicine and CT, NIH/NIBIB, Bethesda, USA.
3. Department of Clinical Physiology, Nuclear Medicine and PET, Rigshospitalet, University of Copenhagen, Denmark

Corresponding Author:

Silvia Morbelli

Nucler Medicine Unit

IRCCS Ospedale Policlinico San Martino

University of Genoa

Tel: +390105553038

FAX: +390105556911

E-mail: silvia.morbelli@unige.it

**Abstract**

This joint Practice Guideline, or Procedure Standard, was developed collaboratively by the European Association of Nuclear Medicine (EANM) and the Society of Nuclear Medicine and Molecular Imaging (SNMMI). The goal of this guideline is to assist nuclear medicine practitioners in recommending, performing, interpreting, and reporting the results of dopaminergic imaging in parkinsonian syndromes. Currently nuclear medicine investigations can assess both presynaptic and postsynaptic function of dopaminergic synapses. To date both EANM and SNMMI have published procedural guidelines for dopamine transporter imaging with SPECT (in 2009 and 2011, respectively). An EANM guideline for D2 SPECT imaging is also available (2009). Since the publication of these previous guidelines, new lines of evidence have been made available on semi-quantification, harmonization, comparison with normal datasets and longitudinal analyses of dopamine transporter imaging with SPECT. Similarly, details on acquisition protocols and simplified quantification methods are now available for dopamine transporter imaging with PET, including recently developed fluorinated tracers. Finally, [18F]fluorodopa PET is now used in some centers for the differential diagnosis of parkinsonism, although procedural guidelines aiming to define standard procedures for [18F]fluorodopa imaging in this setting are still lacking. All these emerging issues are addressed in the present procedural guidelines for dopaminergic imaging in parkinsonian syndromes.

# Keywords

Brain; DAT; DOPA; PET; SPECT; Parkinson; Parkinsonian Syndromes

# Preamble

The Society of Nuclear Medicine and Molecular Imaging (SNMMI) is an international scientific and professional organization founded in 1954 to promote the science, technology, and practical application of nuclear medicine. Its 15,000 members are physicians, technologists, and scientists specializing in the research and practice of nuclear medicine. In addition to publishing journals, newsletters, and books, the SNMMI also sponsors international meetings and workshops designed to increase the competencies of nuclear medicine practitioners and to promote new advances in the science of nuclear medicine. The European Association of Nuclear Medicine (EANM) is a professional nonprofit medical association that facilitates communication worldwide between individuals pursuing clinical and research excellence in nuclear medicine. The EANM was founded in 1985.

The SNMMI/EANM periodically define new standards/guidelines for nuclear medicine practice to help advance the science of nuclear medicine and to improve the quality of service to patients. Existing standards/guidelines will be reviewed for modifications or renewal, as appropriate, on their fifth anniversary or sooner, if indicated. As of February 2014, the SNMMI guidelines are referred to as procedure standards. Any previous practice guideline or procedure guideline that describes how to perform a procedure is now considered an SNMMI procedure standard.

Each standard/guideline, representing a policy statement by the SNMMI/EANM, has undergone a thorough consensus process in which it has been subjected to extensive review. The SNMMI/EANM recognizes that the safe and effective use of diagnostic nuclear medicine imaging requires specific training, skills, and techniques, as described in each document.

The EANM and SNMMI have written and approved these standards/guidelines to promote the use of nuclear medicine procedures with high quality. These standards/guidelines are intended to assist practitioners in providing appropriate nuclear medicine care for patients. They are not inflexible rules or requirements of practice and are not intended, nor should they be used, to establish a legal standard of care. For these reasons and those set forth below, the SNMMI/EANM cautions against the use of these standards/guidelines in litigation in which the clinical decisions of a practitioner are called into question.

The ultimate judgment regarding the propriety of any specific procedure or course of action must be made by medical professionals considering the unique circumstances of each case. Thus, there is no implication that an approach differing from the standards/guidelines, standing alone, is below the standard of care. To the contrary, a conscientious practitioner may responsibly adopt a course of action different from that set forth in the standards/guidelines when, in the reasonable judgment of the practitioner, such course of action is indicated by the condition of the patient, limitations of available resources, or advances in knowledge or technology subsequent to publication of the standards/guidelines.

The practice of medicine involves not only the science but also the art of dealing with the prevention, diagnosis, alleviation, and treatment of disease. The variety and complexity of human conditions make it impossible to always reach the most appropriate diagnosis or to predict with certainty a particular response to treatment. Therefore, it should be recognized that adherence to these standards/guidelines will not ensure an accurate diagnosis or a successful outcome. All that should be expected is that the practitioner will follow a reasonable course of action based on current knowledge, available resources, and the needs of the patient to deliver effective and safe medical care. The sole purpose of these standards/guidelines is to assist practitioners in achieving this objective.

The present guideline/standard was developed collaboratively by the EANM and SNMMI. It summarizes the views of the Neuroimaging Committee of the EANM and the Brain Imaging Council of the SNMMI, and reflects recommendations for which the EANM and SNMMI cannot be held responsible. The recommendations should be taken into context of good practice of nuclear medicine and do not substitute for national and international legal or regulatory provisions.

# Introduction

Parkinsonian syndromes are a group of diseases that share cardinal signs of parkinsonism, such as bradykinesia, rigidity, tremor, and postural instability. Idiopathic Parkinson’s disease (IPD) is the most common cause of parkinsonism but numerous other etiologies can lead to the presence of parkinsonism. Indeed, parkinsonism can be present in all alpha synucleinopathies which include Lewy body diseases (LBDs) a subset of disorders associated with the accumulation of Lewy bodies (LB) and neurites, i.e., intracytoplasmic inclusions composed of aggregated alpha-synuclein and other proteins such as ubiquitin [1]. The most clinically relevant subtypes of LBDs are IPD and dementia with Lewy bodies (DLB) [1]. Alpha synucleinopathies also include multiple system atrophy (MSA), an atypical parkinsonism characterized by the presence of glial and neural silver staining aggregates of alpha-synuclein [2]. Finally, the tauopathies cortico-basal degeneration (CBD) and progressive supranuclear palsy (PSP) are movement disorders belonging to the spectrum of fronto-temporal degeneration (and they are also defined as atypical parkinsonisms) [3, 4]. The differential diagnosis of the neurodegenerative parkinsonian syndromes includes clinical entities such as essential tremor (ET), drug-induced parkinsonism, vascular parkinsonism, and psychogenic parkinsonism. In patients with ET the tremor is most evident during voluntary movement rather than at rest, although some patients with essential tremor can demonstrate a breakthrough resting tremor and cogwheel rigidity, or other isolated parkinsonian characteristics thus making the clinical diagnosis more challenging [5].

In many patients, the clinical differential diagnosis of parkinsonism is relatively straightforward [1]. However in numerous conditions, an improvement in diagnostic accuracy is possible using dopaminergic imaging [8]. This imaging technology may be particularly helpful in patients with incomplete or atypical syndromes, unsatisfying response to therapy, overlapping symptoms or in patients with early/mildly symptomatic stages of disease.

Currently, nuclear medicine investigations can assess both presynaptic and postsynaptic function of dopaminergic synapses [6-8]. Presynaptic dopaminergic imaging helps clarify the differential diagnosis between neurodegenerative parkinsonian syndromes and non-dopamine deficiency etiologies of parkinsonism [6-8].

Presynaptic dopaminergic function can be summarized as follows. Dopamine is produced via two amino acids. First, L-tyrosine is hydroxylated to form L-dopa, which subsequently is decarboxylated to dopamine by aromatic L-amino-acid decarboxylase (AADC; also known as dopa decarboxylase). Next, dopamine is transported to intracellular vesicles by vesicular monoamine transporter 2 (VMAT2). As a result of neuronal depolarisation, these vesicles are emptied into the synaptic cleft where the synaptic dopamine then interacts with postsynaptic dopamine receptors. After dopamine has been emptied into the synaptic cleft, it can be subject to re-uptake into the presynaptic neuron by dopamine transporter (a.k.a. dopamine active transporter - DAT).

Post-synaptic dopamine receptors can be divided into D1-like receptors (D1, D5) and D2-like receptors (D2, D3 and D4) [9]. Over 90% of D2 receptors are located post-synaptically and so imaging of D2 receptors is frequently referred to as imaging of postsynaptic D2 receptors [10]. To date, three Procedure Guidelines/Procedure Standards have been published by EANM and SNMMI, respectively. Both EANM and SNMMI have published Guidelines/Standards for dopamine transporter imaging with SPECT (in 2009 and 2011, respectively) An EANM Guideline for D2 SPECT imaging is also available (2009) [9, 11-13]. Since the publication of these previous documents, new lines of evidence have been made available on semi-quantification, harmonization, comparison with normal datasets and longitudinal analyses of DAT SPECT. Similarly, details on acquisition protocols and simplified quantification methods are now available for dopamine transporter imaging with PET, including recently developed fluorinated tracers. Finally, especially in some nuclear medicine centers equipped with PET and without cyclotron, [18F]fluorodopa PET may be used for the differential diagnosis of parkinsonism, although procedural GLs aiming to define standard procedures for [18F]fluorodopa imaging in this setting are still lacking.

Goals

The goal of this guideline is to assist nuclear medicine practitioners in recommending, performing, interpreting, and reporting the results of dopaminergic imaging in parkinsonian patients.

# Definitions

1. SPECT - Single Photon Emission Computed Tomography (also known as SPET). This imaging modality allows the visualization of the three dimensional distribution of radiopharmaceuticals labelled with gamma-ray emitting radionuclide such as 123I, 99mTc etc.

SPECT/CT - SPECT imaging combined with CT (Computed Tomography) in a hybrid scanner.

Positron Emission Tomography. This imaging modality is specific for imaging positron-emitting radionuclides such as 11C and 18F through coincidence detection of annihilation photon pairs.

PET/CT - PET imaging combined with CT (computed tomography) in a hybrid scanner. PET may also be combined with magnetic resonance imaging (MRI); however, the use of these scanners in patients with movement disorders has not yet been adequately addressed in the literature.

1. Presynaptic dopaminergic imaging refers to SPECT and PET imaging studies evaluating the integrity of presynaptic nigrostriatal dopaminergic synapses.
2. Post-synaptic dopaminergic imaging refers to SPECT and PET imaging studies that evaluate the integrity of dopaminergic neurons at the post-synaptic level (frequently referred to as imaging of the postsynaptic D2 receptors).

# Clinical indications

Presynaptic dopaminergic imaging is indicated for detecting loss of functional nigrostriatal dopaminergic neuron terminals of patients with clinically uncertain parkinsonian syndromes, especially:

1. To differentiate essential tremor from neurodegenerative parkinsonian syndromes. Note that presynaptic dopaminergic imaging is unable to distinguish IPD and DLB from PSP, CBD, or putaminal variant of MSA [14-17].

2. To help distinguish between DLB and other dementias (in particular, Alzheimer’s Disease, AD) [18-20].

3. To differentiate patients with presynaptic degenerative dopamine deficiency parkinsonism from those with other forms of parkinsonism, e.g. between IPD and drug induced, psychogenic, or vascular parkinsonism [21-23]

4. To detect early presynaptic parkinsonian syndromes [24, 25]

Post-synaptic dopaminergic imaging can help separate typical from atypical parkinsonian syndromes. The main indication is the differentiation of IPD from other neurodegenerative parkinsonian syndromes where loss of D2 receptors occurs (e.g., MSA, PSP) [26]. However, the clinical use of SPECT or PET tracers for post-synaptic dopaminergic imaging is currently limited by several factors and has been in many centers replaced by other molecular imaging targets [27, 28] (see section IX).

Less common indications for post-synaptic dopaminergic imaging in parkinsonian syndromes:

1. Assessment of the extent of D2 receptor blockade during treatment with dopamine D2 antagonists (neuroleptics)

2. Wilson’s disease. Loss of striatal D2 receptor function is related to the severity of neurological symptoms in Wilson’s disease and may show the degree of neuronal damage due to cytotoxic copper deposition in the striatum [29].

# Qualification and responsibilities of the personnel

*Physician*: SPECT and PET examinations should be performed by, or under supervision of a physician specialized in nuclear medicine and certified by accrediting boards. In Europe, the certified nuclear medicine physician who authorizes the study and signs the report is responsible for the procedure, according to national laws and rules. For the United States, see the SNMMI Guideline for General Imaging (Society of Nuclear Medicine., 2010, http://interactive.snm.org/docs/General\_ Imaging\_Version\_6.0.pdf).

*Technologist*: SPECT and PET examinations should be executed by qualified registered/certified Nuclear Medicine Technologists. Please refer to: Performance Responsibility and Guidelines for Nuclear Medicine Technologists 3.1 and <http://www.eanm.org/content-eanm/uploads/2016/11/EANM_2017_TC_Benchmark.pdf> for further details. In some jurisdictions there may be additional qualifications necessary for technologists to also operate CT and/or MR components.

*Physicist:* A certified Medical Physics Expert (MPE) is responsible for quality assurance of SPECT and PET systems that are in clinical use and also for managing possible malfunctions of these systems. The MPE is also responsible for the optimal implementation of procedures considering national and international radiation protection safety standards, both for patients and for personnel.

# Procedure/specifications of the examination

**Request/history**

1. The nuclear medicine imaging facility should check with the radiopharmaceutical provider to ensure availability before scheduling the exam. Advanced notice may be required for tracer delivery.

2. The requisition should include a brief description of symptoms and the clinical question. Information should be obtained regarding the following:

a. Past or current recreational drug use, head trauma, stroke, psychiatric illness, epilepsy, or tumor.

b. Neurologic symptoms: type, duration, and left- or right-sidedness.

c. Current medications and when last taken.

d. Patient’s ability to lie still for approximately 30–45 min "(for SPECT imaging with [123I]FP-CIT).

e. Prior brain imaging studies (e.g., CT, MRI, PET, and SPECT).

**Radiopharmaceuticals**

*Presynaptic tracers*

– [123I]β-CIT: 2β-carboxymethoxy-3β-(4-[123I]iodophenyl)tropane.

–[123I]FP-CIT: N-3-fluoropropyl-2β-carbomethoxy-3β-(4-[123I]iodophenyl)nortropane ([123I]-ioflupane)

– [99mTc]TRODAT-1: [2[[2-[[[3-(4-chlorophenyl)-8-methyl-8-azabicyclo[3,2,1]-oct-2-yl]-methyl](2-mercaptoethyl)amino]ethyl]amino]ethanethiolato(3-)-N2,N2’,S2,S2]oxo-[1R-exo-exo)])- [99mTc]-technetium

– [18F]fluorodopa: 3,4-dihydroxy-6-[18F]fluoro-L-phenylalanine;

– [11C]-PE2I) N‐(3‐iodopro‐2*E*‐enyl)‐2β‐carbo[11C]methoxy‐3β‐(4′‐methylphenyl)nortropane

[18F]-FE-PE2I: N-(3-iodoprop-2E-enyl)-2β-carbo[18F]fluoroethoxy-3β-(4-methylphenyl)-nortropane

-(+)[11C]dihydrotetrabenazine: (+)2-Hydroxy-3-isobutyl-9-[11C]methoxy-10-methoxy-1,2,3,4,6,7,-hexahydro-11bH-bezo[α]-quinolizine; (+)[11C]DTBZ

-(+)[18F]FP-DTBZ: (+)-α-9-O-(3-[18F]fluoropropyl)DTBZ; [18F]AV-133

*Post-synaptic tracers*

– [123I]IBZM: (S)-3-[123I]iodo-N-[(1-ethyl-2-pyrrolidinyl)]methyl-2-hydroxy-6-methoxybenzamide

– [123I]epidepride: ( S)-N-((1-ethyl-2-pyrrolidinyl) methyl)-5-[123I]iodo-2,3-dimethoxybenzamide; [123I]epidepride is the iodine analogue of isoremoxipride (FLB 457)

– [18F]fallypride: 5-(3-[18F]fluoropropyl)-2,3-dimethoxy-N-[(2S)-1-prop-2-enylpyrrolidin-2-yl]methyl]benzamide; [18F]N-allyl-5-fluorpropylepidepride

– [18F]desmethoxyfallypride: (S)-N-((1-allyl-2-pyrrolidinyl)methyl)-5-(3-[18F]fluoropropyl)-2-methoxybenzamide; [18F]DMFP

– [11C]raclopride: (2S)-3,5-Dichloro-N-[(1-ethyl-2-pyrrolidinyl)methyl]- 6-hydroxy-2-[11C]methoxybenzamide

# Procedure/specifications of the examination for presynaptic dopaminergic imaging with SPECT using 123I-labelled dopamine transporter ligands and with PET using [18F]fluorodopa

As of early 2019, among SPECT tracers only [123I]FP-CIT has been approved by both the U.S. Food and Drug Administration (FDA) and the European Medicines Agency (EMA) for testing of dopaminergic neuronal integrity in suspected parkinsonian syndromes. A clinical alternative for DAT SPECT imaging is [18F]fluorodopa PET. [18F]fluorodopa is approved by EMA but is not currently approved by U.S. FDA for marketing by commercial radiopharmacies. In the U.S. [18F]fluorodopa can be administered to human subjects under an Investigational New Drug application (IND), approved by the FDA for clinical research purposes. From the pathophysiological point of view, it should be noted that the effect of dopaminergic neurodegeneration as detected with [18F]fluorodopa PET tends to be smaller than with DAT SPECT [30]. Indeed the pathophysiology of DAT and DOPA reduction in the nigrostriatal degeneration follows different timeframes as the two processes are parallel but not synchronous. This is presumably due to the presence, in the presymptomatic and early symptomatic phases, of opposite compensatory mechanisms to the reduction of the number of dopaminergic terminals: on one hand the reduction of DAT, which increases the synaptic availability of dopamine, and on the other hand the upregulation of DOPA conversion to dopamine by aromatic L-amino-acid decarboxylase (AADC) in the nerve terminal [31]. A recent, meta-analysis of 142 positron emission tomography and single photon emission computed tomography studies demonstrated that AADC defect is consistently smaller than the dopamine transporter and vesicular monoamine transporter 2 defects, suggesting upregulation of AADC function in IPD [30,31]. However, available studies in small number of patients and controls demonstrated that DAT SPECT and [18F]fluorodopa PET scans are both able to diagnose presynaptic dopaminergic deficits in early phases of PD with excellent sensitivity and specificity [32]

**SPECT using iodine-123 labelled dopamine transporter ligands**

## Patient preparation and precautions

*Prearrival*

Drugs that bind to the dopamine transporter with high affinity or mobilize endogenous dopamine may interfere with DAT binding ligands. Whether or not discontinuation of these drugs prior to tracer administration may minimize the interference has not been determined by controlled *in vivo* or *in vitro* studies. The benefits and risks of stopping medication that may interfere with the reliability of the information in the images should be evaluated on an individual basis. Nevertheless, in general, it is recommended that prior to the investigation, patients should avoid taking any medication or drugs of abuse that could significantly affect visual and semi-quantitative analysis of DAT binding ligands (except if the specific aim of the study is to assess the effect of such medication on DAT binding) [33]. See Table 1.

1. Check for medications or drugs that may alter tracer binding, and (if possible) stop such medication for at least 5 half-lives.
2. Cocaine, amphetamines, and methylphenidate are high affinity DAT blockers that severely decrease tracer binding to DAT. Modafinil, used in the treatment of narcolepsy, at therapeutic doses significantly blocks dopamine transporter binding similarly to methylphenidate [34]. Drugs that elevate endogenous dopamine levels should be discontinued. Indeed, while not psychoactive, cannabidiol acts as an anandamide reuptake inhibitor which can decrease DAT function [35].
3. Ephedrine and phentermine, particularly when used as tablets, some antidepressant (bupropion, radafaxine), anesthetics (ketamine, phencyclidine, and isoflurane), opioids (fentanyl) and anticholinergic drugs (benzatropine) may also decrease, to a lesser extent, DAT ligands binding
4. Selective serotonin reuptake inhibitors (SSRI) may increase both [123I]β-CIT and [123I]-FP-CIT striatal binding to DAT by 20 % and 10%, respectively [36]. The underlying mechanism is unknown. It has been speculated that mild [123I]-FP-CIT blocking of the SERT in the occipital cortex might result in a higher striatal binding ratio, although direct interaction between the dopaminergic and serotonergic system cannot be excluded [37]. Sertraline, has a relative higher affinity for DAT that might lead to decreased striatal tracer binding. However, it has been hypothesized that these effects may be counterbalanced by the 10% increase in binding related to serotonin blocking [37]. Overall, these effects, although small (especially for [123I]-FP-CIT), should be considered in the framework of research studies but they should not significantly affect the interpretation of visual assessments in clinical routine application [33].
5. Chronic lithium treatment might significantly reduce DAT binding as demonstrated in a patient with clinical parkinsonism, a finding reversed after lithium withdrawal in parallel with recovery from parkinsonian symptoms. A possible mechanism, could be the impaired synaptic DA release induced by lithium leading to compensatory down-regulation of membrane DAT [37].
6. Cholinesterase inhibitors and neuroleptics probably do not interfere significantly with tracer binding to DAT [36].
7. Anti-Parkinsonian drugs (e.g., L-DOPA, dopamine agonists, monoamine oxidase B inhibitors, N-methyl-D-aspartate receptor blockers, amantadine, and catechol-O-methyltransferase inhibitors in standard dosages) do not interfere with [123I]-FP-CIT binding to DAT to any significant degree [36, 38].
8. It should be noted that nicotine use (smoking, vaping, chewing tobacco) may interfere with DAT availability [39]; however, such an effect would be too small to lead to misinterpretation of an individual scan.

An extensive overview of drug effects on DAT SPECT can be found in Booij and Kemp [33]

*Preinjection*

To reduce exposure of the thyroid to free 123I, a single dose of potassium iodide oral solution or Lugol's solution (equivalent to 100 mg iodide), potassium perchlorate (400 mg), or sodium perchlorate (600 mg) should be administered at least one hour before administration of [123I]-FP-CIT. The use of any of these products should be avoided in patients with known sensitivities. Even in the absence of a blocking agent, the radiation dose to the thyroid would be low [13].

*Precautions*

Continuous supervision of the patients during the entire scanning procedure is necessary. This is especially important for patients with uncooperative patients due to cognitive/behavioural symptoms that may be present in some parkinsonian patients.

*Preparation of the radiopharmaceutical*

Radiopharmaceuticals will be delivered ready to use.

*Quality control check*

Radiochemical purity and other quality control parameters as listed in the package inserts need to be checked. The instructions provided by the manufacturer should be followed.

## Injection

Inject the radiopharmaceutical intravenously as a slow bolus over approximately 20 second followed by a saline flush of the intravenous line. Inject the radiopharmaceutical within the time frame given by the manufacturer (generally on the day of delivery).

*Administered Dosage*

1. Adults: 110–250 MBq (typically 185 MBq) for both [123I]FP-CIT and [123I]β-CIT
2. Children: Currently, there are no established clinical indications in children. Accordingly, its safety and efficacy have not been established in pediatric patients. If indicated in the future, activity should be administered according to the recommendations of the EANM Pediatric Task Group.
3. The effect of renal or hepatic impairment on DAT SPECT imaging has not been established. Because [123I]-FP-CIT is excreted by the kidneys, patients with severe renal impairment may have increased radiation exposure and altered DAT SPECT images. The patient should also be encouraged to stay hydrated and void frequently the remainder of the day to minimize radiation exposure to the bladder [14]. Hypersensitivity and injection site reactions have been reported. In clinical trials, the most common adverse reactions were headache, nausea, vertigo, dry mouth, and dizziness which occurred in less than 1% of subjects.

C. Equipment Specifications

1. Multiple detector (triple or dual head) or other dedicated SPECT cameras for brain imaging should be used for data acquisition [40, 41]. Single detector units cannot generally be recommended as prohibitively long scan times is needed to acquire the required number of counts. They may be used only if the scan time is prolonged appropriately, injected activity in the upper permissible range is used, and meticulous care is taken to produce high-quality images.
2. Low Energy High (or Ultra high) resolution (LEHR or LEUHR) parallel-hole collimators are the most frequently available collimator sets for brain imaging. It should be noted that, if available, fan-beam collimators have a more advantageous trade-off between resolution and count rate capability with respect to parallel-hole collimators. All-purpose collimators are not suitable. The use of medium energy collimators could be advantageous regarding septal penetration; however, they are hampered by a lower spatial resolution [42]. If available, collimator sets specifically adapted to the characteristics of 123I may be used. [42]

## D. Acquisition protocol

*Timing*

1. SPECT should be started when the ratio of striatal-to-occipital tracer binding is stable [21]. In fact, waiting for stability of the ratio guarantees the most reliable data from a semi-quantitative point of view and is optimal for visual assessment when low background interference is needed.
	1. [123I]FP-CIT: 3 to 6 h after injection
	2. [123I]β-CIT: 18 to 24 h after injection.
2. It is recommended that each center maintains a standardized post-injection imaging time (e.g., always start imaging at 3 hours post radiotracer injection for [123I]FP-CIT) to ensure that data are comparable between subjects and intra-individual follow-up studies.
3. Since tracer binding is not determined by flow effects, patients do not need to stay in a quiet/dim environment. There are no dietary restrictions, pre- or post-injection.
4. If the patient is claustrophobic, or is unable to lie still, mild sedation with a short-acting benzodiazepine may be given prior to imaging and will not affect scan quality or quantification. In this circumstance, the patient should not operate machines or drive a car for the rest of the day

*Positioning*

1. The patient should void immediately prior to imaging.
2. The patient should lie on the couch in a supine position with the head resting in the head holder in a comfortable orientation. A light head restraint, such as a strip of tape across the chin or forehead, may be considered to minimize motion artefacts.
3. Although proper alignment of the head would be preferable, patient comfort is more important in order to prevent movement during the scan. Tilt and orientation can be adjusted during post-processing.
4. The arms should be supported at the side or in front of the patient. Additional leg support may be used to improve comfort. Binding the shoulders (e.g., with a sheet) may also help prevent movement and reduce the orbital radius of the camera heads.

The striatum (the caudate nucleus and putamen) and occipital cortex must be in the field of view. The inclusion of cerebellum might be omitted in the acquisition to optimize radius of rotation and spatial resolution only in patients with challenging positioning of the head [43]. It should be noted, however, that when the cerebellum is not in the field of view, reorientation accuracy and spatial registration might be affected (see below).

*Acquisition*

Acquisition parameters:

1. Rotational radius: smallest possible with appropriate patient safeguards (typically 11-15cm).
2. Photopeak: 159 KeV ± 10% or 159 KeV+/- 7.5 %, as recommended by manufacturer. Additional energy windows for scatter correction may be used as needed.
3. Mode: step and shoot mode is used predominantly. Continuous mode acquisition may provide shorter total scan times and reduce mechanical wear to the system
4. Matrix: at least 128x128
5. Zoom: adjust zoom factor to achieve a pixel size of 2.5-4 mm. In general acquisition pixel size should be one-third to one-half of the expected resolution.
6. Rotation: 360° coverage
7. Angular step: ~3° increments (for step and shoot mode).
8. Frame time: the number of seconds per position depends on the sensitivity of the system, but usually 25–40s are required.
9. Total detected events in the photopeak window: should be > 1.5 million total counts for [123I]-FP-CIT; >1 million for [123I]β-CIT. In this framework, it is helpful, prior to start images acquisition, to use lateral or anterior projections to ensure that the count rate is appropriate to achieve the required number of counts.
10. SPECT/CT: the SPECT image acquisition on SPECT/CT systems should follow the general recommendations for DAT SPECT image acquisition.
11. CT-based attenuation correction (CTAC) may be performed with SPECT/CT systems following manufacturer’s protocol. Overall in the few available studies comparing data obtained using CTAC with data obtained using uniform AC, minor differences were observed, without impact on diagnostic performance. It should be noted that when SPECT/CT devices are available, CT is generally used for attenuation correction and anatomical localization only. Accordingly, the CT acquisition mode should use the protocol that delivers the lowest possible radiation dose to the subject (e.g., a relatively low dose protocol) that retains the quantitative accuracy of corrections for attenuation and anatomic localization. According to Guidelines of the Quantitative Imaging Biomarkers Alliance of the Radiological Society of North America (QIBA-RSNA), when CT is used for attenuation correction only, the CT can be performed with 5 – 10 mAs. [44].
12. Segmentation of data acquisition into multiple sequential acquisitions may permit exclusion of data with artefacts, e.g., remove segments of projection data with patient motion.

## Image processing

*Review of projection data*

* 1. Prior to reconstruction, the raw projection images should be watched in cine mode or in sinogram mode to check for movement, which may be difficult to recognize in the reconstructed SPECT slices.
	2. Given the 8-10 mm spatial resolution of SPECT, small amounts of motion may be tolerated. For more significant motion, repeated scanning is recommended.

*Image reconstruction*

1. Methods: filtered back projection (FBP) or iterative reconstruction. In general, FBP appears to be more consistent across vendors. However, iterative reconstruction may be preferred to help reduce streak artefact and for more robust physical correction algorithms.
2. If iterative reconstruction is used, typically, about 100 EM equivalent iterations (iterations x subsets) should be used, e.g., 10 iterations and 10 subsets for 120 projections or 12 iterations and 8 subsets for 128 projections [45]. However, if normal databases for semiquantitative assessment are used, given that the choice of iterations and subsets will affect semiquantitative values, chosen reconstruction parameters should match those used in the normal databases.
3. Ideally the entire brain volume should be reconstructed to assist in the reformatting of data. Semiquantitative software may, in fact, require the cerebellum as a reference region or the inclusion of the cerebellum may be important in the spatial registration process.

*Filtering*

1. Data should be filtered in all three dimensions (x,y,z). This can be achieved either by two-dimensional prefiltering of the projection data or by applying a three-dimensional postfilter to the reconstructed data. The second option is recommended for iterative methods, because of the underlying assumption of Poisson distribution of the detected counts.
2. Low-pass (e.g., Butterworth or Gaussian) filters should generally be used to reduce image noise, but not at the detriment of spatial resolution. Spatially-varying filters should be used with caution, as they may produce artefacts.
3. If resolution modelling is incorporated into the reconstruction, the noise suppression introduced in the algorithm may remove or reduce the need for additional filtering.

*Corrections*

1. The use of corrections (attenuation, scatter and septal penetration- referred in the following simply as “scatter correction”, and partial volume effect) do not necessarily benefit visual interpretation although it will significantly affect the values derived from semiquantitative analysis. Corrections may be considered to reduce quantitative bias and, when accuracy is desired, all three corrections should be applied [46, 47].
2. If normal databases are to be used, the reconstruction should match that used to create the normal database.
3. Chang attenuation correction with a uniform attenuation map can be used effectively with an appropriate linear attenuation correction coefficient for 123I and careful contouring of the head/scalp. Typically, the linear attenuation factor should be in the range μ=0.10 -0.12 cm-1 without scatter correction, while the μ value in water of 0.143 cm-1 should be used with scatter correction [9, 11, 12].
4. CT based attenuation correction is more robust and can account for head holder attenuation. Obviously, the choice of low-dose ‘attenuation correction only’ CT versus high-dose ‘attenuation correction and diagnostic’ CT has important implications for the radiation dose, and a full diagnostic CT should be used only when justified by the clinical indication.
5. Scatter (and septal penetration) correction can improve visual contrast and reduce bias in semi-quantitative analysis. It can be performed using either energy-window based corrections [48] or transmission-dependent convolution subtraction [49]. The former is simple to implement and can account for extra-cerebral activity but increases image noise and therefore worsens quantitative precision. The latter handles noise better, but is unable to account for septal penetration of 123I high energy gammas, which can vary significantly across scanners [50]. Modern vendor software allows for the application of scatter and attenuation correction and adherence to manufacturer-specified parameters is recommended
6. Due to the limited spatial resolution of SPECT systems, imaging small structures in the basal ganglia suffers from large partial volume effects (PVE), which causes the most significant bias/loss in signal during semi-quantitative analysis [51]. Resolution modelling techniques in the reconstruction can compensate for some of these losses, although overshoot artefacts associated with the Gibbs phenomena [60] mean that such techniques should be used with caution. It should also be noted that for some commercial systems, application of resolution modelling is combined with other corrections and cannot be decoupled.

## F. Quality Control

*Prerequisites*

1. All SPECT systems should be subjected to a quality control program (AAPM (In press), EANM 2010, IAEA 2009) [52,53]
2. For SPECT/CT systems, CT sub-systems should also be under a quality control program (IAEA 2012) and have additional SPECT/CT QC tests performed [53]
3. The SPECT system must be set up and calibrated for imaging 123I.
4. Given the small structures in the basal ganglia, reconstructed SPECT spatial resolution should be ≤ 10 mm (FWHM).
5. If using Chang attenuation correction, 123I SPECT uniformity using the reconstruction and corrections for clinical imaging should be checked to ensure that an appropriate linear attenuation factor is being used.
6. It can be helpful to image an anthropomorphic striatal phantom to assess SPECT image quality using clinical acquisition and reconstruction parameters.

*Periodic Quality Control (QC)*

1. Intrinsic (without the collimator) uniformity to 123I should be checked.
2. Centre of rotation checks with the collimators used for DAT SPECT should be assessed.
3. If using SPECT/CT systems, the alignment/registration of SPECT and CT should be checked to ensure good localization and attenuation correction.

*QC on the day of DAT SPECT*

1. Uniformity (intrinsic or system) of detector response should be checked daily prior to patient injection.
2. The collimators should be checked for mechanical damage. If damage is detected or suspected, perform a high-count extrinsic uniformity test.
3. The peaking of the camera to 123I should be checked with the patient on the scanner.

## Interpretation/Semiquantification

*Reformatted sections and reorientation:*

Transaxial slices should be reformatted into three orthogonal planes for visual assessment. Generate transverse sections parallel to a given anatomic orientation (e.g. anterior commissural to posterior commissural line; AC-PC) ensuring a high degree of standardization in plane orientation. Create coronal and sagittal sections orthogonal to the transverse sections.

*Visual interpretation*

Because patients usually do not become symptomatic before a substantial number of striatal synapses have degenerated, visual interpretation of the scan is usually sufficient for clinical evaluation in the majority of patients. Several studies show excellent results with trained readers using visual interpretation only [15-18, 54, 55].

1. Images should be read on the computer screen rather than from hard copy, because this allows variation in color table and adjustments of background subtraction or contrast. Readers are recommended to select one color scale for consistency. Level of contrast and background subtraction should be carefully checked as inappropriate thresholding may result in artefacts.
2. Data evaluation should consider age and relevant morphological information (CT, MRI). Specific attention should be paid to known structural lesions in the basal ganglia, midbrain and to the structures selected as reference region for semiquantitative evaluation. Review of any available CT head scans or MRI brain scans may give additional information. Known anatomic lesions and central atrophy may alter the tracer binding, location or shape of the striatal structures. A direct co-registration or side-by-side reading of an equivocal scan with an MRI scan may assist in excluding or estimating vascular comorbidity.
3. Visual assessment is usually sufficient for evaluating striatal left/right symmetry and striatal subregions.
4. On transaxial images the normal striata should look crescent- or comma-shaped and should have symmetric well-delineated borders, although mild asymmetry may occur in normal subjects.
5. Care should be taken in the alignment of the head, particularly in the coronal plane, in order to avoid artifactual left-right asymmetry which may lead to visual misinterpretation of the scan. It is recommended, therefore, to visualize all the transaxial slices containing striatal signal so that real asymmetry can be differentiated from an insufficiently corrected head tilt.
6. Abnormal striata will have reduced intensity on one or both sides, often shrinking to an oval or circular shape. The level of striatal activity should be compared with the background activity. Orthogonal slices should be used. Some decrease in striatal binding, in both the caudate and putamen, occurs with normal aging (see below paragraph on semiquantification).
7. Common patterns for abnormalities emerge on visual interpretation: in IPD, the disease first becomes visible in the dorsal putamen contralateral to the neurologic signs [54, 55]. In fact, in early IPD, the putamen is usually more severely affected than the caudate nucleus. Furthermore the putamen of the less affected hemisphere is also early involved (often before the caudate nucleus of the most affected side).This pattern results in the so-called dot-shape (as opposed to the normal comma-shape) typical of IPD in the earliest stages of disease.
8. The presence of a dot-shape with uptake reduction in both putamina has been repeatedly described in both patients with hemiparkinsonism and even in patients with pre-motor PD (i.e. in clinical trials carried out in patients with anosmia and/or REM-sleep behavior disorders [24, 25, 56, 57]. These findings testify to the high sensitivity of DAT SPECT in the earliest stages of disease and its high negative predictive value which still remains one of its main strengths.
9. In IPD the DAT binding is often first reduced in the putamen contralateral to the side of the body showing the most severe signs of disease, then the reduction in tracer uptake progresses to the contralateral putamen and anteriorly and ipsilaterally to the caudate over time. Correlation of scores/scales measuring severity of motor impairment and disability (UPDRIII and Hoehn and Yahr Scale) on one side and degree of DAT binding as assessed by DAT SPECT on the other have been demonstrated [58]. Accordingly, IPD usually is characterized by reduced uptake in all the basal ganglia structures only in advanced stages of disease. A presence of strong decrease in signal-to noise-ratio in patients with very mild motor symptoms might be due to tracer extravasation.

In DLB, early involvement of the caudate nucleus may cause the difference in tracer uptake between the putamen and the caudate to be less apparent than in PD. This may result in images with "weak commas" or "balanced loss" [59]. In general, atypical parkinsonian syndromes such tend to be more symmetric and to involve the caudate nucleus relatively earlier as shown in group analyses. However, there is too much overlap between the disease patterns to allow for adequate discrimination between PD/DLB, MSA, PSP, and CBD on an individual level [18, 59].

1. Since late 90’s, some patients (10% to 15 %) that were enrolled in PD clinical trials as having newly diagnosed PD, demonstrated normal DAT SPECT and were originally defined as subjects having a scan without evidence of dopaminergic deficit (SWEDD) [60, 61]. The concept of SWEDD is, however, highly debated and controversial and represents a heterogenous entity presumably due to misdiagnosis and/or SPECT misinterpretation. Longitudinal assessment demonstrated that subjects with a SWEDD at baseline continue to have a SWEDD in follow-up scans. In fact, it was more recently demonstrated that they have minimal evidence of clinical or imaging PD progression. This new evidence strongly suggests that SWEDD subjects are unlikely to have IPD [60, 61].
2. In essential tremor, psychogenic parkinsonism and drug-induced parkinsonism, DAT ligands binding is normal allowing an accurate differential diagnosis with respect to neurodegenerative parkinsonian syndromes [22, 23, 55].
3. The term vascular parkinsonism is controversial and ill-defined. Basal ganglia vascular lesions are very common and will cause parkinsonism in only a minority of patients, and neither infarct site nor the size can predict the clinical presentation. In the literature, DAT ligands binding has been described as normal or only slightly diminished, except when an infarct directly involves a striatal structure. Even then, a deficit from an infarct often gives a “punched-out” appearance, differing in morphology and quality from a typical presynaptic parkinsonian syndrome deficit. If clarification is needed, a recent MRI scan should be reviewed [62-64]. However, an infarct placed adjacently to striatal structures may give a more diffuse reduction through the PVE mimicking early IPD. It should be noted that, vascular lesions in or near the basal ganglia are common in the typical age group scanned and is a rare cause of parkinsonism. Enlargement of perivascular spaces (Virchow–Robin spaces) are generally most prominent in the inferior basal ganglia, and may lead to a diffuse reduction in striatal uptake following the distribution of spaces [65]. As “incidental” cerebrovascular lesions are found in 20-30% of patients with histologically confirmed IPD, the basal ganglia structural lesions are not uncommonly combined with a true nigro-striatal dysfunction [66].
4. DAT SPECT differentiates between Alzheimer’s disease (AD) and DLB with a high degree of accuracy [20]. In fact, striatal binding is usually normal or only mildly diminished in AD and is significantly decreased in DLB [19-20]. In particular, given its high specificity (20% higher with respect to clinical diagnosis), the revised criteria for the diagnosis of DLB further reinforced the role of DAT SPECT considering it as a diagnostic biomarker [20]. It should be noted, however that approximately 10% of patients who met pathologic criteria for DLB can show a normal DAT SPECT at the time of the clinical diagnosis (possibly becoming abnormal after 1.5-2 years) [67, 68]. In a recent autopsy study, this finding was related to topographical heterogeneity of α-synuclein deposition [19]. In fact this subgroup of patients was characterized by only mild cell loss and α-synuclein pathology of low severity in the substantia nigra and moderate severity at the cortical level.
5. Consistent with the clinical and neuro-radiological asymmetry characterizing CBD, a more asymmetric reduced uptake (involving both putamen and caudate on the hemisphere contralateral to the most affected body side) has been reported in patients with CBD [69]. DAT SPECT is not sufficiently accurate to support the differential diagnosis between IPD and CBD. However, a markedly asymmetrical pattern of DAT binding reduction should be specifically emphasized in the final report if the clinical suspicion is CBD in order to guide the subsequent diagnostic workout (i.e. 123I-mIBG and 18FDG PET might be useful in this scenario) [70, 71]. Unexpectedly, normal DAT SPECT findings might also occur in patients with clinical corticobasal syndrome (CBS). This finding is likely to be related to an underlying etiology different from CBD (AD, fronto-temporal dementia) [72]. However this issue needs further clarification that might be provided by further studies based on multi-modal biomarker (CSF, amyloid PET and possibly TAU PET) in patients with CBS/CBD.
6. DAT SPECT is abnormal in 30-60% of patients with sporadic fronto-temporal dementia (FTD), but usually less pronounced than in DLB/IPD [72-75]. However, some cases with marked abnormal DAT SPECT have been reported in specific subtypes of FTD (i.e. familial FTD with parkinsonism linked to chromosome 17q (FTDP-17) [72-75]. DAT SPECT is not a direct marker of α-synuclein pathology but only a measure of its effect on nigrostriatal neurons, which can be damaged by other pathology [76].

 A schematic summary of expected DAT SPECT findings in the most frequent clinical entities included in the differential diagnosis of parkinsonism is provided in Table 2.

*Semi-quantification*

The output images from image reconstruction are considered the input for image analysis.

1. In the clinical setting, semiquantitative analysis can be an adjunct to visual interpretation. Some studies have shown the addition of quantitative information results in improved diagnostic performance [50, 77, 78]. In research settings, semiquantitative methods are considered to provide a more precise assessment of dopamine transporter density, allowing for more reproducible measurements of disease and response to therapy.
2. The most commonly used semi-quantitative outcome measure in both research and clinical settings is the specific binding ratio (SBR) calculated as the striatal target-to-background ratio [79, 80]. SBR is related to, but not a specific measure of the density of dopamine transporters; however, is not a direct measure of DAT density (Bmax), synaptic density, or neuronal number. Many factors which are unrelated to the density of dopamine transporters influence the determination of SBR (Tables 3 reports both biological and technical factors affecting SBR).
3. Even though semiquantification seems straightforward, there can be considerable inter-observer variation and errors in the placement of the regions of interest (ROIs), or in the reorientation of the brain, which can lead to variability in the interpretation. This variability may be reduced using automated systems analyzing volumes of interest. Furthermore, for interpretation, semiquantitative data must be compared with a suitable database of reference values, preferably age-matched.
4. Potential advantages of semiquantification include more objective measurements of SBR (with improved inter-reader agreement and reader confidence) and the ability to obtain a quantitative result that can be correlated with loss of presynaptic dopaminergic neurons [81]. In addition, if a reference database of age-matched reference values is available, other potential advantages include earlier disease detection, more accurate interpretations in patients who are difficult to classify visually, higher interrater agreement and greater usefulness for research studies
5. ROI-based techniques can be used to assess specific DAT binding in the striatum and striatal subregions (caudate nucleus, putamen). If manually drawn ROIs are used, usually transverse (axial) slices are selected for ROIs definition. For semiquantitative evaluation, either only the slices with the highest striatal binding are chosen or the entire striatal volume might be taken into account. The use of standardized ROIs is helpful in limiting operator-variability to the positioning task. The shape of the template ROIs could be either geometrical or anatomical, the size should be at least twice the FWHM.
6. Reference regions with absent (or low) DAT density are used to assess nonspecific binding. The reference region is ideally the cerebellum, as it contains no known dopaminergic neurons. The occipital cortex can be also used particularly when the axial field of view is limited. The methodology used for defining the reference region should be consistent across all patients measured. It is particularly important when comparing measurements at different time points in the same patient.
7. If anatomical scans (MRI or CT) are available, volumetric regions encompassing the anatomic extent of the basal ganglia can be used. This is particularly important when low specific binding is expected (e.g. in case of a severe loss or blockade of the DAT), especially in the absence of an automated method.
8. Once the Volumetric ROI (VOI) are placed, SBR values are generally obtained as:

(Mean Counts of striatal VOI – Mean Counts of background VOI)/(Mean Counts of the background VOI)

1. Automated-semiquantification

Several approaches have been proposed to perform semiquantification of DAT SPECT, and several commercial and freeware software are available.

Some of them include a one or two step-based fully automated registration of the patient SPECT scan to a template or to an averaged, spatially normalized brain volume. Predefined VOIs are then automatically placed on the striatal and the background regions [82-88]

 VOIs can also be defined for striatal substructures, allowing to quantify the SBR for caudate and putamen, as well as other parameters such as asymmetry between left and right putamen and caudate, putamen-to-caudate ratio (PCR), and caudate-putamen ratio (CPR).

The availability of these parameters can be useful in complex or borderline cases. In particular PCR is most sensitive in borderline and early stage cases, because of a higher degree of independence on camera, reconstruction algorithm, and background region. Some studies have reported a mild reduction of PCR with age, but there are conflicting data in the literature on this topic (possibly due to the number and age range of included subjects, and the modalities of VOI drawing) [87]. Evaluation of asymmetry between left and right sides is also useful in the earliest stages, although mild asymmetry between striata or striatal subregions may occur in normal subjects [88].

All the numeric parameters are dependent on acquisition and reconstruction parameters and on the corrections applied [49, 50]. As a result, there are no universal (age-dependent) cutoffs for normal vs. abnormal [89-91]. Accordingly, each site would ideally need to establish its own reference range by scanning a population of healthy controls or needs to calibrate its procedure with the characteristics of normal controls databases (see below).

Differences in SBR also occur as the result of VOI strategy, with two approaches commonly employed. The first incorporates regions for the caudate and putamen on left and right sides which tries to capture the anatomic bounds of these structures. The second approach utilizes small regions of interest sampling the tracer uptake in the caudate, mid-putamen, and posterior putamen on the right and left sides [92, 93].

A different approach for determining the SBR is used in the so-called Southampton method [93], based on the measurement of total counts from each striatum. It uses striatal VOIs of geometrical shape, sufficiently large to capture all counts originating from the striatum including those detected outside its anatomical boundaries, and a background region derived from the whole brain minus the striatal VOIs. The SBR is then calculated as follows:

 VolStrVOI/VolStr\*(CountsStrVOI/counts per voxelbkg \* VolStrVOI -1)

Individual computed tomography (CT) or MRI-guided methods [94] and, more recently, machine-learning techniques for pattern recognition as well as parametrization of textural patterns have also been proposed as user-independent methods for DAT SPECT results classification. These methods still remain in the realm of promising research [95, 96].

Some studies have applied PVE in the process of binding computation of the caudate nucleus, putamen and background, but the added value of PVE in this settings has not been systematically investigated and is not recommended in routine clinical practice.

1. Available healthy subjects’ database

As previously mentioned, great caution must be exercised when using SBR data from the literature as differences regarding to the camera system how the images were acquired, processed, and analyzed will result in different values [50, 87, 97, 98].

Calibrations across cameras have been described using anthropomorphic phantoms to develop relative correction factors for standardization between instruments [46, 87, 99]. The QIBA of the RSNA is developing a protocol for quantitative standardization of [123I]-FP-CIT SPECT analysis which will detail a validated protocol for imaging centers to apply in order to obtain the high reproducibility [45].

In recent years, two normative database have been provided:

 The ENC-DAT (European Normal Control Database of DaTSCAN) database was developed from 2009. The study was promoted by the EANM Neuroimaging Committee, thanks to a cooperative effort among different imaging centers. The purpose of the multicenter project was to collect a large number of [123I]FP-CIT SPECT scans of healthy controls and to provide reference images and values of DAT availability measures obtained across a wide age range and with balanced gender representation. The final database included SPECT data from 139 healthy controls (74 men, 65 women; age range 20–83 years, mean 53 years) acquired in 13 different centers [100].

The second database of healthy controls was collected in the framework of the Parkinson’s Progression Marker Initiative (PPMI) which is a multicenter, international, longitudinal study evaluating clinical, biochemical, and imaging measures of PD progression. Features of PPMI include; 1) incorporation of DAT SPECT to confirm the presence or absence of a DAT deficit for enrolled PD and healthy volunteers, respectively, 2) rigorous standardized acquisition protocols, and 3) central core lab reconstruction of raw projection data for subsequent uniform analyses to be made available to the investigator community. PPMI includes 423 progressing Parkinson’s patients and 196 age-matched controls studied over 3 to 4 years with serial [123I]-FP-CIT SPECT imaging and other clinical, imaging, and fluid biomarkers [101].

Normative data provided by these two multicenter initiatives demonstrated that normal aging is associated with about 5.5-6% signal loss per decade (0.6%/y). Gender had also an effect on SBR in the caudate and putamen (often not considered in the software available for comparison) [81, 94, 96]. Dependency of DAT on BMI, handedness, circadian rhythm or season was not demonstrated [87, 100, 102].

SPECT/CT: Very few studies tested the use of CT images acquired on the SPECT/CT system to derive either warping parameters for VOI or for regional SBR assessment, and the anatomical information provided by CT did not demonstrate a relevant impact on diagnostic performance [103].

1. Longitudinal data

Serial imaging within the individual patient in order to track disease progression by visual or semiquantitative analysis might be clinically useful only once it has been demonstrated that signal size is able to capture the slow progression of the disease process.

Besides providing a healthy subjects database, the PPMI initiative has provided an assessment of within-subject changes of SPECT imaging in a PD patient cohort followed longitudinally over 3-5 years. The aim is to understand the utility of DAT SPECT as a putative biomarker of PD progression after early motor signs of PD appear [104]. In fact if intraindividual comparison is performed (i.e. baseline vs. follow-up for therapy control or assessment of disease progression), more subtle changes might need to be highlighted.

In the PPMI study, regional SBR were measured in ipsilateral and contralateral caudate, anterior putamen, and posterior putamen at baseline and 1, 2, and 4 years during the follow up of IPD patients. During the 4 years, there was a significant longitudinal change in dopamine transporter binding in all striatal regions [104]. Available analyses seem to suggest that given the size of signal change in an early PD cohort, [123I]FP-CIT SPECT might be a suitable biomarker of PD progression. In fact, initial longitudinal data suggest SBR reductions over one year are approximately 20 times the rate of signal loss seen in normal aging [104].

H. Multi-center trial harmonization

In order to harmonize imaging equipment sites for multicenter research or trials there are several areas to be considered [105].

1. Data acquisition: the acquisition protocol should be harmonized as much as possible. Collimators, energy windows, pixel sizes and required number of counts are key considerations. During this process, attention should be given to manufacturer specific requirements e.g. energy windows for scatter correction.
2. Image reconstruction and filtering: The choice of reconstruction and filtering is very important in harmonization particularly given that semiquantification is strongly dependent on the approach taken and the corrections included. Variability in reconstruction algorithms and functionality across centers may drive the choice of reconstruction to an approach that is comparable and achievable by centers, i.e. typically one with no corrections. Alternatively, to overcome this variability, reconstruction at a central core lab using a generic algorithm could be considered.
3. Comparability of system performance: the use of an anthropomorphic striatal phantom can help assess comparability of visual and semiquantitative performance across sites/scanners and is, therefore, recommended [46]. One or more data acquisitions with varying filling ratios is suggested to determine the relationship between scanner semiquantitative performance over a range of SBRs.

**Presynaptic dopaminergic imaging with PET using [18F]fluorodopa**

## Preparation of the radiopharmaceuticals

[18F]fluorodopa as all radiopharmaceuticals must be produced by qualified personnel according to cGMP, or using GMP-compliant methods that conform to regulatory requirements. If the radiopharmaceutical is delivered ready to use, the quality control is carried out by the manufacturer prior to delivery of the final product.

1. Patient preparation and precautions

*Prearrival*

1. Anti-parkinsonian drugs (L-dihydroxyphenylalanine (L-DOPA), dopamine agonists, monoamine oxidase B inhibitors, N-methyl-D-aspartate receptor blockers, and amantadine) do not significantly interfere with visual interpretation. However a discontinuation of L-DOPA by at least 12 hours should be considered as it has been recommended by manufacturers and applied in published studies to improve image quality (although it should not interfere visual interpretation significantly). Neuroleptics have a limited effect on [18F]fluorodopa kinetics and do not need to be discontinued [106].
2. Food intake affects [18F]fluorodopa kinetics, since dietary large neutral amino acids compete with [18F]fluorodopa for transfer by a common carrier through the blood brain barrier [107]. Therefore, it is recommended that patients should not intake any amino-acid containing foods 4 h prior to the procedure.
3. Smoking probably does not interfere significantly with fluorodopa kinetics (although dopamine synthesis capacity seems to be increased in heavy smokers) [108].
4. Recent use of recreational drugs that directly act on the dopaminergic system (cocaine, amphetamine) should be noted.

*Preinjection*

1. Prehydration is important to reduce the radiation exposure of [18F]fluorodopa by stimulating bladder voiding. Patients should be encouraged to drink sufficient amounts of water and to empty their bladder prior to and after the PET examination.
2. Pretreatment with the peripheral dopa decarboxylase inhibitor carbidopa (150mg or 2mg/kg up to 150mg, orally) or catechol-O-methyltransferase (COMT) inhibitor entacapone (200 mg, orally) 60-90 minutes prior to [18F]fluorodopa injection is a common practice that increases the availability of DOPA to the brain and reduces absorbed dose to the bladder and kidneys [109-111].

## Injection

The recommended injected activity for brain imaging in adults for PET-CT is 185 MBq. If no carbidopa is given ahead of [18F]fluorodopa higher doses may be necessary. [18F]fluorodopa should be injected as a bolus.

## Image Acquisition

Positioning

The scan should be performed with the patient in supine position, with the head in a dedicated holder, and arms along the body. The entire brain should be in the field of view, including the entire cerebellum. Avoid extreme neck extension or flexion.

1. Head stability

The patient should be informed immediately before PET acquisition to avoid head movements during all parts of the investigation

Head stability can be obtained by comfortably positioning the patient in the head holder. The head can be further secured by using flexible restraints such as tape or padding."

During the entire investigation, continuous visual/video monitoring of the patient is necessary.

*Acquisition protocol for PET/CT systems [112, 113]*

1. The preferred sequence of PET imaging is:

CT scout topogram for PET(/CT) to setup field of view

For attenuation correction: Low-dose CT scan (see before paragraphs on SPECT/CT), or transmission scan. If a PET/MRI system is used, an MR-AC scan for MR-based attenuation correction should be acquired.

Static or dynamic single Field of View (FOV) PET acquisition

1. Image acquisition should be performed in 3D data acquisition mode, and attenuation correction should be based on low-dose CT or on a 511keV-transmission scan. If 511keV-transmission scanning is used, the transmission images should be acquired before tracer injection. CT parameters should always be chosen to ensure the lowest radiation exposure to the patient that is compatible with this purpose.
2. A fixed time for the start of image acquisition should be taken.
3. Static acquisition: 10-20 min static image acquisition obtained 70-90 minutes post-injection is recommended [114, 115]. If movement artifacts are expected, it can be helpful to acquire the static time window dynamically e.g. in 5-min frames, or in list-mode, check the sinograms, and use only the sinograms of the properly acquired motion-free time period for reconstruction.
4. Dynamic or list-mode acquisition [115, 116]: a protocol of at least 90 minutes acquisition is required for the calculation of the [18F]fluorodopa influx constant (Ki). The acquisition starts at the time of injection using short 30 second time frames that progressively increase to 5 minutes in duration. During the first 10 minutes post injection, the following image acquisition frames can be used: 6 frames of 30 seconds and 2 frames of 210 seconds allowing to precisely obtain information of the tracer uptake phase. During the remaining 80 minutes, frames of 5 minutes (or even 10 minutes) are used for the evaluation of [18F]fluorodopa conversion to 18F-fluorodopamine by aromatic amino acid decarboxylase and uptake and trapping of 18F-fluorodopamine into synaptic vesicles.

## Image reconstruction

* 1. Iterative reconstructions are currently the standard. When Ordered-Subsets Expectation Maximization (OSEM) type methods are used the number of iterations depends on the equipment characteristics, the injected activity, the acquisition duration as well as on the corrections included in the reconstruction process.
	2. Matrix sizes and zoom factors during reconstruction should consider the small size of the caudate and putamen. When possible, matrix sizes and zoom factors during reconstruction should be large enough in order that reconstructed voxel sizes are within 2.0 – 3.0 mm in any direction.
	3. Regular corrections are necessary before or during image reconstruction such as attenuation, scatter, random, dead time and decay corrections and sensitivity normalization. To date, the benefit of time-of-flight correction has been modest, but it can be applied [117, 118].
	4. Resolution modelling during reconstruction, called point-spread function (PSF) reconstruction deserves specific consideration. Resolution modelling has been developed to compensate for cross-contamination between adjacent functional regions with distinct activities, referred to as PVE [119]. In theory, given the size of the striatal structures, it would be beneficial to reconstruct the exact activity distribution using this correction. However, there is a limit to the recovery achievable by PSF correction, even at high iterations due to the loss of high frequency information during data acquisition [119]. The application of PSF correction modifies the noise structure and it produces a “lumpiness” aspect [120] (referred to as the Gibbs phenomenon). Therefore, PSF reconstruction artefacts can lead to misinterpretations when used to analyse small subcentimeter structures and is currently not recommended for striatal imaging
	5. Spatial filters applied during or after reconstruction should be selected as not to impair the spatial resolution needed for striatal imaging. However, the choice is highly dependent on the acquisition data (noise) and the entire reconstruction process. Ideally the final reconstructed spatial resolution should aim at a FWHM < 6 mm.
	6. Head-movement correction can be performed if the acquisition is performed in list mode or multi frame setting (dynamic scans). Data acquired during patient movement can be discarded before reconstruction or corrected post reconstruction with dynamic or list mode acquisitions.

##  Interpretation/Semiquantification

The uptake of [18F]fluorodopaover the first 90 minutes into the striatum primarily reflects influx and decarboxylation of the tracer to [18F]fluorodopamine [121].

*Visual Analysis*

General concepts already discussed for visual analysis and interpretation of [123]FP-CIT SPECT also apply to [18F]fluorodopa PET images. See below further details and corresponding references specifically related to [18F]fluorodopa.

1. Images should be read on the computer screen, since it is of paramount importance to adjust the color scale appropriately. Indeed, the first step of visual reading is the setting of the maximum color scale value. The reader should look for the voxel of maximal value within the striatum and select this value as the maximum of the chosen color scale. The position of this maximal uptake is important for the interpretation (see below).
2. Care should be taken in the 3D alignment of images to insure the proper appreciation of real uptake asymmetry. It is recommended to visualize all the transaxial slices containing striatal uptake as well as those including midbrain in order to be able to detect extra-striatal uptake.
3. Normal [18F]fluorodopa striatal uptake is much less dependent on age than DAT tracers.
4. Data evaluation should take into account relevant morphological information (CT, MRI). Specific attention should be paid to known structural lesions in the basal ganglia. A direct co-registration or side-by-side reading of an equivocal scan with an MRI scan may assist in excluding or estimating vascular comorbidity.
5. In normal subjects, the striatal uptake is homogeneous within the striatum and symmetrical. It is important to know that the local maximum of striatal uptake is normally located in the putamen. Extra-striatal uptake of [18F]fluorodopa due to monoaminergic innervation is low and normally undetectable.
6. The presence of an asymmetric pattern of reduced putamen and preserved caudate uptake is visually characteristic of PD, the low uptake level in the most affected putamen accurately separates IPD cases from healthy controls in published series [122]. Putamen uptake reduction starts from the its posterior part creating a posterior-anterior gradient. It is very useful to identify the local maximum of striatal uptake which is located in the caudate in that case.
7. In case of nigrostriatal degeneration, the mid-brain extra-striatal uptake becomes more visible due to global reduction of striatal [18F]fluorodopa uptake. Catecholaminergic (locus ceruleus) and dopaminergic (substantia nigra) nuclei may demonstrate a “Mickey Mouse” shape in the mid-brain [123].
8. In most of the cases, atypical parkinsonian syndromes cannot be separated from IPD. However, it is noteworthy, that in CBD and in PSP, the caudate sparing which is characteristic of IPD is often lost. In cases of vascular origin, localized defects corresponding to micro-infarctions can be detectable. Direct registration side-by-side reading with MRI scan help interpretation by showing the spatial correspondence of MR striatal lacunae and localized [18F]fluorodopa defects.
9. In essential tremor, psychogenic parkinsonism and drug-induced parkinsonism, [18F]fluorodopa uptake is normal.

*Semiquantification*

1. A simple approach for quantitating tracer uptake is to compare the striatal signal to the signal in the occipital (SOR) or cerebellar cortex where little decarboxylation occurs [122]. This approach has been shown to sensitively detect the asymmetric putamen [18F]fluorodopareductions present in IPD and the SOR reductions correlate with ratings of disability.
2. When a dynamic acquisition is performed, it is possible to compute [18F]fluorodopa influx constants (Ki) from the brain time-activity curves (TACs). In fact Parametric images of specific [18F]fluorodopa influx constants (Ki maps) can be created at a voxel level for the whole brain using linear graphical analysis of TACs with an occipital cortex non-specific reference input function [124, 125]. The Ki is the product of the free 18F-DOPA brain volume of distribution (Vd) at equilibrium, and the decarboxylation rate constant k3 [126, 127].Basically, K1 is thought to reflect [18F]fluorodopa transport from arterial plasma over the blood brain barrier into the dopaminergic neurons and the decarboxylation of [18F]fluorodopa to [18F]-Dopamine [126, 127]. Ki is the product of the free 18F-DOPA brain volume of distribution (Vd) at equilibrium and the decarboxylation rate constant k3 [127]. Ideally, this graphical approach requires a metabolite corrected plasma input function. Nevertheless, a reference region like occipital cortex (or cerebellum) where only minimal amounts of specific [18F]fluorodopa metabolites are assumed to exist (due negligible decarboxylation) can be used as a measure of plasma free [18F]fluorodopa signal avoiding blood sampling [128]. The non specific activity concentration is further assumed to be identical in the striatum and in the occipital cortex. The reference tissue derived influx constant Ki from brain uptake over 30-90 minutes then becomes the product of the [18F]fluorodopa decarboxylation rate constant k3 and the ratio of free brain [18F]fluorodopa / ([18F]fluorodopa +[18F]fluoromethyldopa) Vd.
3. In practice, Ki derived using the reference region approach and simple ratios of striatal / occipital fluorine-18 uptake at 90 minutes are highly correlated, and both approaches differentiate early PD from healthy controls. Either approach can be used for quantitation although only few studies have to-date simplified Ki method with full kinetic analysis.
4. An advantage of the Ki over the ratio approach is that it provides a more direct measure of the [18F]fluorodopa decarboxylation rate constant.
5. Part of [18F]fluorodopamine is subsequently metabolised to [18F]FHVA (6-[18F]fluorohomovanilic acid) and [18F]FDOPAC (3,4-hidroxy-6-[18F]fluorophenylacetic acid). Accordingly, a dynamic PET scan at 3-4 hours after [18F]fluorodopa administration reflects signal washout after 90 minutes due to combined MAOB and central COMT activity. PET can, therefore be used to measure dopamine turnover in the striatum as a ratio of signal 18F loss (kloss) to influx Ki [128]. This approach, while potentially more sensitive than isolated Ki or SOR ratio measurements, requires scans at 90 minutes and again 3-4 hours for the subject so is currently a research rather than a clinical tool [129].

## Equipment Specifications

* 1. State-of-the-art 3D PET/CT or PET/MRI systems should preferably be used. The equipment should allow for the acquisition of data needed for attenuation and scatter correction. Low dose CT is the preferred option, although PET/MRI can use MRI sequences dedicated to attenuation and scatter correction of PET emission data. Transmission sources are also capable of producing adequate attenuation maps on dedicated brain PET only systems.
	2. The equipment should have an axial field of view >15 cm to assure sufficient coverage of the entire brain, including the cerebellum and brain stem.
	3. The PET camera should be able to acquire both static and dynamic, frame or list mode PET emission data in 3D mode.

## Quality control and Interinstitutional PET system performance harmonization

The present guidelines are focused on the use of dopaminergic imaging in parkinsonian syndromes. However, there are both technical and imaging physics related uncertainties that apply to any PET examination regardless of radiotracer or specific application. These aspects have been detailed in previous guidelines [130]. These recommendations should be considered for the use of brain PET examinations in multicenter studies and/or when data are compared to a reference database or disease patterns. In order to guarantee sufficient image quality, quantitative performance, and image harmonization, the correct performance of the PET system must be regularly checked by several QC experiments that have also already previously listed in EANM and joint EANM/SNMMI guidelines. Cross-calibration of the PET(/CT) system against the locally used dose calibrator to prepare and measure patient-specific radiotracer activities is needed. Cross-calibrations should be performed following EARL recommendations and criteria (http://earl.eanm.org).

**VII. Documentation and reporting for presynaptic dopaminergic imaging with either SPECT using iodine-123-labelled dopamine transporter ligands or with PET using [18F]fluorodopa**

**Reporting**

Reports should include all pertinent information, including the name of the patient and other identifiers, such as date of birth, name of the referring physician(s), type and date of examination and patient history including the reason for requesting the study.

1. Body of the report

*Procedures and materials*

1. Include a brief description of the imaging procedure: radiopharmaceutical including administered activity, scan acquisition delay and assessment of scan quality (if compromised give the reason, e.g. motion artefacts, etc.).
2. If sedation is performed, briefly describe the procedure including type and time of medication given in relation to the radiotracer injection. (at least in theory, halogen based anesthetics and benzodiazepines can increase endogenous dopamine levels opening the potential for false positives).

*Findings*

1. Describe the subjective visual impression of striatal binding compared with background activity. Examine both the caudate nuclei and the putamina for decreased activity; note which regions, if any, appear decreased. Note any significant asymmetries; mild asymmetry may occur in healthy individuals. If abnormalities are present, report the location and intensity of the areas of decreased activity. Comment on signal to noise ratio (visualization of extra-striatal uptake or not). State what criteria were used for interpretation (visual assessment or semiquantitative measures, comparison with normal database, etc.).
2. If semiquantitative analysis for DAT SPECT is performed, report the values. If a normal database for patients’ value comparison, report reference range and/or z-score differences with respect to the database in case of abnormal findings (see discussion of utility and limits of database of healthy subjects in previous sections).
3. Correlate the findings with relevant anatomic changes displayed on any available diagnostic CT or MRI scans
4. Limitations: where appropriate, identify factors that could have limited the sensitivity and specificity of the result of the examination (i.e. movement or concomitant medication). State whether the patient used interfering drugs, and if so, which drugs.
5. Clinical issues: The report should address or answer any pertinent clinical issues raised in the request for the imaging examination.

*Comparative data*

Comparisons with previous examinations and reports, if available, should be part of the report indicated. However, in view of the high sensitivity and early impairment characterizing presynaptic dopaminergic imaging, a follow up scan is rarely needed or indicated.

1. Interpretation and conclusions

*Conclusions*

The conclusion should state whether a presynaptic dopaminergic deficit is present or absent. Abnormal findings indicate a presynaptic striatal dopaminergic deficit, which is consistent with a variety of diagnoses, including IPD, PSP, MSA, CBD and DLB. However, when caudate sparing characteristic of IPD is lost, diagnoses of CBD or PSP can be suggested, especially with [18F]fluorodopa which allows higher spatial resolution. Normal findings could suggest essential tremor, drug-induced parkinsonism, psychogenic parkinsonism, AD, or a state of health. For properly selected patients within the approved indication, abnormal findings would be consistent with tremor due to presynaptic parkinsonian syndromes rather than essential tremor. To aid the referring clinician, descriptors such as mild, moderate, or severe can be used to characterize any deficits. A more comprehensive diagnostic comment taking advantage of the available results of other previously performed molecular (i.e. ([18F]FDG PET or [123I]mIBG) or structural imaging studies (i.e. MRI with a punched out appearance of the striatum) could be included in selected cases. Similarly, when appropriate, follow-up or additional studies ([18F]FDG PET, perfusion SPECT, cardiac [123I]mIBG scintigraphy or D2 imaging) can be recommended to clarify or confirm the suspected diagnosis.

*Sources of error and Artifacts*

1. Artifacts (patient movement, camera-related, induced by inappropriate processing; tracer extravasations).
2. Interference from medications (although interactions should affect tracer uptake the whole striatum and should not influence the presence of an asymmetrical pattern of reduced uptake).

# New Tracers for Dopaminergic Imaging in Parkinsonian Syndromes

**PET imaging of DAT using 11C-PE2I and 18F-FE-PE2I**

Among several DAT radioligands available for PET imaging, [11C]-PE2I (N‐(3‐iodopro‐2*E*‐enyl)‐2β‐carbo[11C]methoxy‐3β‐(4′‐methylphenyl)nortropane) was developed in early 2000 in order to image and quantify the DAT not only in the striatum, but also in extrastriatal regions, such as the substantia nigra [131-133]. [11C]PE2I is highly selective for the DAT, displays high target-to-background ratio and good test-retest reliability [134]. It has been already used in the context of clinical trials and PET studies in patients with PD and in animal models of parkinsonism [135-139].

Despite the suitable properties as a PET radioligand for the DAT, absolute quantification requires long duration of imaging for at least 90 min, due to slow wash-out from the brain [140]. In addition, a radioactive metabolite of [11C]PE2I has been found to cross the blood-brain barrier and accumulate in the striatum in rodents. Therefore, further efforts have been made in order to develop an analog of [11C]PE2I with improved imaging properties. [18F]FE-PE2I (N-(3-iodoprop-2E-enyl)-2β-carbo[18F]fluoroethoxy-3β-(4-methylphenyl)-nortropane)is a fluoro-ethyl analog of [11C]PE2I developed at Karolinska Institute and highly selective for the DAT with improved pharmacokinetic properties compared with [11C]PE2I: i) faster washout from the brain and ii) lower production of a radiometabolite (< 2% of plasma radioactivity) that crosses the blood-brain barrier [141-144].

The quantification can be done using the cerebellum as reference region [144]. A simplified method for the quantification of the SBR can also be used by acquiring a static image around the peak specific binding between 16.5 and 42 min after injection [145].

The affinity of [18F]FE-PE2I for the DAT (*K*i=12 nM) and the high target-to-background ratio permit to quantify the DAT in the striatum and in the substantia nigra [143-146]. Compared with other tropane analogs including ß-CIT and FP-CIT, FE-PE2I is by far more selective for the DAT than for other monoamine transporters [131].

Initial studies have shown that in early PD, DAT is decreased by approximately 70% in the putamen and 30% in the substantia nigra, indicating a preferential loss of the DAT in the axonal terminals [146].

The dosimetry of [18F]FE-PE2I has been estimated in non-human primates and human subjects [147, 148]. The estimated effective dose (ED) is 0.023 mSv/MBq, which is similar to that of other 18F-labeled radioligands for brain imaging [149].

Further studies with [18F]FE-PE2I are ongoing to examine the DAT loss across different stages of IPD, the test-retest reliability, clinical most appropriate imaging interval and the longitudinal assessment in IPD patients. Head-to-head comparative studies with [123I]FP-CIT are also ongoing to examine the relative performance of [18F]FE-PE2I in differentiating IPD from controls. The first published comparative study reported that binding potential (*BP*ND) of [18F]FE-PE2I was highly correlated with [123I]FP-CIT data, with intraclass correlation coefficients >0.9 [149]. Similar effect sizes were reported for both radioligands, indicating that [18F]FE-PE2I has similar capability to differentiate patients with parkinsonism from controls as [123I]FP-CIT [149].

 Further studies (and probably long term follow up) will be needed to assess if patients with uncertain borderline [18F]FE-PE2I PET that cannot be explained by structural change/lesions can be better separated based on high or low uptake in the substantia nigra.

The primary clinical benefits compared to DAT SPECT include the improved resolution of PET, an improved clinical workflow, patient and caregiver convenience as a thyroid-blocking agent is not required, and shorter injection-to-scan and acquisition times.

Radiopharmaceutical´s characteristics

1. Administered activity to adults: 100–250 MBq (typically 200 MBq)
2. PET data acquisition:

Suggested dynamic acquisition for estimation of binding potential (*BP*ND): duration, 60 min

Suggested static acquisitions: for best simplified estimation of specific binding ratios: 25 min from 17 to 42 min.; For routine clinical use: 10 min from 30 to 40 min

1. Image processing:

Measured or CT attenuation correction

OSEM reconstruction- Resolution/PSF modelling not needed for routine clinical use

Final image resolution: 5-8 mm

1. Quantification

Binding potential (*BP*ND) estimated with simplified reference tissue method or Logan graphical analysis using cerebellum as reference region [144].

1. Semiquantification

SBR calculated with cerebellar grey matter as reference. Target metrics: SBR Caudate, SBR Putamen, SBR Striatum; Putamen/Caudate ratio; Hemisphere asymmetry index = (R−L)/(R+L)

**PET imaging of vesicular monoamine transporter type-2 (VMAT2)**

After synthesis, monoaminergic neurotransmitters are concentrated in synaptic vesicles for exocytotic release. The vesicular monoamine transporter type-2 (VMAT2) is expressed by all monoaminergic neurons and serves to transport transmitter from cytoplasm into vesicles [150]. In the CNS, VMAT2 is expressed exclusively by monoaminergic (dopaminergic, serotonergic, norepinephrinergic, or histaminergic) neurons [151], albeit it is also expressed in pancreatic beta cells and several monoaminergic neurons of the human enteric nervous system [152]. Over 95% of striatal VMAT2 binding sites are associated with dopaminergic terminals [153]. In vivo imaging and quantification of VMAT2 has been reported for a series of radioligands based on the vesicle-depleting drug tetrabenazine. The most widely used ligand for research purposes is (+)-[11C]dihydrotetrabenazine ([11C]DTBZ), which binds specifically and reversibly to VMAT2 and is amenable to quantification of striatal, diencephalic and brainstem neurons and terminals with PET [154, 155]. As only the (+)-enantiomer binds specifically to VMAT2, use of the racemic tracer reduces specificity with ensuing lower image contrast. Despite the successful application of (+)-[11C]DTBZ for human studies, clinical utilization will likely require radioligands with longer half-lives. Fortunately, development of fluorine-18-labeled VMAT2 radioligands such as (+)-[18F]FP-DTBZ ((+)-α-9-O-(3-[18F]fluoropropyl)DTBZ) provides a 18F-labeled VMAT2 ligand that will be more suitable for routine use in clinical practice [156]. When compared with (+)-[11C]DTBZ, it is not only more favorable due to the longer half-life fluorine-18 labelling, but also has higher affinity for the VMAT2 receptor (Ki=0,3 nM vs 0,1 nM).

Radiopharmaceutical´s characteristics

1.Administered activity to adults: 300-450 MBq.

2.Radiation Dosimetry:

Radioactivity uptake in the brain is the highest at 7.5 ± 0.6% injected dose at 10 min after injection. High absorbed doses were found in pancreas, liver and upper large intestine wall. The highest-dosed organ, which received 153 ± 24 μGy/MBq, was the pancreas. The effective dose for (+)-[18F]FP-DTBZ was 28 ± 3 μSv/MBq. These values are comparable with those reported for other fluorine-18 labeled radiopharmaceuticals [156].

3. PET data acquisition:

A single 10-min PET scan can be acquired 90 min post-injection in 3D mode. Scanning time of 90-100 min for (+)-[18F]FP-DTBZ is considered as the optimal time window for summed uptake measurements in terms of correlation of standardized uptake value ratio (SUVRs of striatum to Occipital Cortex) to distribution volume ratios (obtained with dynamic acquisition) as well as in terms of differential power, stability and clinical feasibility across and between patients with IPD and control subjects [157].

4. Image Processing:

Measured or CT attenuation correction

PET images can be reconstructed using 3D OSEM algorithm

5.Visual assessment:

(+)-[18F]FP-DTBZ binding in normal control subjects shows symmetric and highest uptake of (+)-[18F]FP-DTBZ in striatum, followed by nucleus accumbens, hypothalamus, substantia nigra, and raphe nuclei [158].

6.Quantification analysis:

Striatal VMAT2 density of anterior putamen is higher than posterior putamen, which is higher than that of the caudate nucleus. Lowest uptake is seen in the cortex with essentially no specific binding in the occipital cortex [158].

# Procedure/specifications of the examination for postsynaptic dopaminergic imaging

Historically, the most widely applied radiotracers for imaging D2-like receptors with SPECT has been [123I]IBZM [159, 160]. For PET, [11C]raclopride, [18F]fallypride and [18F]Desmethoxyfallypride (DMFP) have been used. These dopamine receptor antagonist derivatives are not selective radiopharmaceuticals for the D2 receptor, since they also bind to the D3 receptor [160]. However, the vast majority of D2-like receptors in the striatum are D2 receptors. EANM GLs aiming to describe standard procedures for brain neurotransmission using dopamine D2 receptor ligands have been published in 2009 [9]. These guidelines dealt with indications, assessment, processing, interpretation and reporting of D2 imaging. The procedural recommendations defined in these GLs are still presently valid with the exception of details on equipment specification which should be updated in agreement of what already described in previous paragraphs of the present GLs. It should be noted that the main added clinical value of D2-imaging in patients with parkinsonian syndromes was the proposed capability of this imaging to support the differential diagnosis within neurodegenerative parkinsonisms by differentiating between diseases with (i.e. atypical parkinsonism such as MSA, CBD and PSP) and without (i.e. IPD) degeneration at post-synaptic level [10]. However D2 SPECT imaging is at the moment not clinically available in several countries and it has been suggested that FDG PET outperforms D2 SPECT Imaging for the differential diagnosis between PD/DLB and the other atypical parkinsonisms [27]. Similarly [123I]mIBG myocardial scintigraphy can be used to differentiate between IPD from MSA, PSP and CBD [161]. Accordingly updated data on the use of these tracers are beyond the aims of the present clinical use-oriented standard procedures and the procedural published by the EANM NeuroImaging Committee in 2009 should be still considered as the main reference for D2 dopaminergic imaging in parkinsonian syndromes [9].

# Radiation safety

In this section we present a set of values and advice that are compliant with the legal requirements in the European Basic Safety Standards Directive and with IAEA Safety Guides and the SNMMI Guideline for General Imaging [162, 163]. However, the way the EU-directive is being implemented in the national member states in Europe varies considerably. In addition, regional and local rules, e.g. from local ethical committees, may also apply in a specific setting and supersede this guideline. The general system of radiation protection is based on the three concepts of 1) Justification 2) Optimization and 3) Dose limits.

Justification for an examination at the individual level is implicit when following the indications in this guideline. The set of acquisition parameters given in this guide serves to optimize the activity used (and the dose provided) for current standard equipment. Further local optimization may apply when special, more sensitive equipment is available. Dose limits do not apply for medical exposure of patients; protection is managed through the justification and optimization only. For biomedical research performed on patients or volunteers, however, ICRP (ICRP62 and ICRP103), WHO and EC (EC RP99) has issued guidance balancing the value for society of the expected outcome of the research against (effective) dose. This is the source of information that ethics committees (institutional review boards) will typically consider.

In addition – since medical use of radiation is a “planned exposure” - the concept of dose constraints applies to “comforters and cares”, to relatives, to “other” hospital staff not working directly with ionizing radiation, and to members of the public. The relevant radiation exposure for these groups is highly dependent on local factors and procedures. It therefore falls outside the scope of this guideline and must be handled by the associated medical physics expert.

For the tracers described in this guideline, Effective Dose and Absorbed Dose (per MBq) to the highest exposed organs has been compiled from the available literature and is given in Table 4 [40, 148, 156, 164-170]. In most cases, following this guideline will lead to effective doses below 5 mSv per examination.

If a CT scan is utilized for attenuation correction only, the lowest possible exposure settings should be used, resulting in a contribution to effective dose of less than 0.1 mSv.

In radiation protection, there is a special concern for a fetus (pregnancy) and small children (breastfeeding) due to their higher radiosensitivity. General rules for obtaining information about pregnancy status must be established and followed as locally implemented. Neither pregnancy nor breastfeeding are absolute contraindications for imaging using ionizing radiation, but require careful considerations for the justification. The need to support the diagnosis of parkinsonian syndromes with SPECT or PET examinations rarely apply to pregnant or breastfeeding patients.

However, like for any other diagnostic procedure in a female patient known or suspected to be pregnant, a clinical decision is necessary in which the benefits are weighed against the possible (hypothetical) harm. If the procedure consists of a static scan it may be possible to reduce the activity (hence the absorbed dose to organs and effective dose) and prolong the acquisition time to maintain image quality; a reduction in image quality is not recommended.

Breastfeeding**:** Before administering radiopharmaceuticals to a mother who is breastfeeding consideration should be given to the possibility of delaying the examination (administration of radionuclide) until the mother has ceased breastfeeding and to what is the most appropriate radiopharmaceutical choice, bearing in mind the possible secretion of activity in breast milk. Unfortunately, since these events are rare, the data available about this secretion is also sparse and often uncertain. Formally (in the terminology of radiation protection) the child is a “member of the public”, and therefore, the protection of the child is based on a dose constraint of 1 mSv (ICRP103, BSS, IAEA). This dose constraint is the combined result of an intake of radioactivity and a contact dose. Ideally these two sources should be considered separately, since contact dose might be eliminated if the mother expresses milk that can be given to the child by a third person. The recent IAEA Safety Guide, however, only gives recommendations that include both sources and are obviously different for tracers labelled with 123I, 11C or 18F (crf IAEA Safety Guide, Appendix III).

# References

1. Berardelli A, Wenning GK, Antonini A, Berg D, Bloem BR, Bonifati V, Brooks D, Burn DJ, Colosimo C, Fanciulli A, Ferreira J, Gasser T, Grandas F, Kanovsky P, Kostic V, Kulisevsky J, Oertel W, Poewe W, Reese JP, Relja M, Ruzicka E, Schrag A, Seppi K, Taba P, Vidailhet M. EFNS/MDS-ES/ENS [corrected] recommendations for the diagnosis of Parkinson's disease. Eur J Neurol. 2013 Jan;20(1):16-34

2. Martí, M.J.; Tolosa, E.; Campdelacreu, J. Clinical overview of the synucleinopathies. Mov Disord 2003, 18 Suppl 6, S21-S27.

3. Armstrong MJ, Litvan I, Lang AE, Bak TH, Bhatia KP, Borroni B, Boxer AL, Dickson DW, Grossman M, Hallett M, Josephs KA, Kertesz A, Lee SE, Miller BL, Reich SG, Riley DE, Tolosa E, Tröster AI, Vidailhet M, Weiner WJ. Criteria for the diagnosis of corticobasal degeneration. Neurology. 2013 Jan 29;80(5):496-503

4. Höglinger GU, Respondek G, Stamelou M, Kurz C, Josephs KA, Lang AE, Mollenhauer B, Müller U, Nilsson C, Whitwell JL, Arzberger T, Englund E, Gelpi E, Giese A, Irwin DJ, Meissner WG, Pantelyat A, Rajput A, van Swieten JC, Troakes C, Antonini A, Bhatia KP, Bordelon Y, Compta Y, Corvol JC, Colosimo C, Dickson DW, Dodel R, Ferguson L, Grossman M, Kassubek J, Krismer F, Levin J, Lorenzl S, Morris HR, Nestor P, Oertel WH, Poewe W, Rabinovici G, Rowe JB, Schellenberg GD, Seppi K, van Eimeren T, Wenning GK, Boxer AL, Golbe LI, Litvan I; Movement Disorder Society-endorsed PSP Study Group. Clinical diagnosis of progressive supranuclear palsy: The movement disorder society criteria. Mov Disord. 2017 Jun;32(6):853-864

5. Hughes AJ, Ben-Shlomo Y, Daniel SE, Lees AJ. What features improve the accuracy of clinical diagnosis in Parkinson’s disease: a clinicopathologic study. Neurology. 2001;57(suppl 3):S34–S38.

6. Elsinga PH, Hatano K, Ishiwata K. PET tracers for imaging of the dopaminergic system. Curr Med Chem. 2006;13:2139–53

7. Thobo is S, Jahanshahi M, Pinto S, Frackowiak R, Limousin Dowsey P. PET and SPECT functional imaging studies in Parkinsonian syndromes: from the lesion to its consequences. Neuroimage. 2004;23:1–16.

8. Booij J, Knol RJ. SPECT imaging of the dopaminergic system in (Premotor) Parkinson’s Disease . Parkinsonism Relat Disorders 2007;13(Suppl 3):S425–8.

9. Van Laere K, Varrone A, Booij J, Vander Borght T, Nobili F, Kapucu OL, Walker Z, Någren K, Tatsch K, Darcourt J. EANM procedure guidelines for brain neurotransmission SPECT/PET using dopamine D2 receptor ligands, version 2. Eur J Nucl Med Mol Imaging. 2010 Feb;37(2):434-42.

10. Kim YJ, Ichise M, Ballinger JR, et al. Combination of dopamine transporter and D2 receptor SPECT in the diagnostic evaluation of PD, MSA, and PSP. Mov Disord. 2002;17:303–12.

11. Darcourt J, Booij J, Tatsch K, Varrone A, Vander Borght T, Kapucu OL, Någren K, Nobili F, Walker Z, Van Laere K. EANM procedure guidelines for brain neurotransmission SPECT using (123)I-labelled dopamine transporter ligands, version 2. Eur J Nucl Med Mol Imaging. 2010 Feb;37(2):443-50

12. Djang DS, Janssen MJ, Bohnen N, Booij J, Henderson TA, Herholz K, Minoshima S, Rowe CC, Sabri O, Seibyl J, Van Berckel BN, Wanner M. SNM practice guideline for dopamine transporter imaging with 123I-ioflupane SPECT 1.0. J Nucl Med. 2012 Jan;53(1):154-63.

13. DaTSCAN European Public Assessment Report. 072200en6.pdf and DaTSCAN-H-266-AR-II.pdf, http://www.emea.europa.eu/humandocs/Humans/EPAR/datscan/Datscan.htm

14. Benamer TS, Patterson J, Grosset DG, Booij J, de Bruin K, van Royen E, et al. Accurate differentiation of parkinsonism and essential tremor using visual assessment of [123I]-FP-CIT SPECT imaging: the [123I]-FP-CIT study group. Mov Disord 2000;15 (3):503–10.

15. Booij J, Hemelaar TG, Speelman JD, de Bruin K, Janssen AG, van Royen EA. One-day protocol for imaging of the nigrostriatal dopaminergic pathway in Parkinson’s disease by [123I]FPCIT SPECT. J Nucl Med 1999;40(5):753–61.

16. Scherfler C, Schwarz J, Antonini A, Grosset D, Valldeoriola F, Marek K, et al. Role of DAT-SPECT in the diagnostic work up of parkinsonism. Mov Disord 2007;22(9):1229–38.

17. Nicastro N, Garibotto V, Burkhard PR. 123I-FP-CIT SPECT Accurately Distinguishes Parkinsonian From Cerebellar Variant of Multiple System Atrophy. Clin Nucl Med. 2018 Feb;43(2):e33-e36.

18. McKeith I, O’Brien J, Walker Z, Tatsch K, Booij J, Darcourt J, et al. Sensitivity and specificity of dopamine transporter imaging with 123I-FP-CIT SPECT in dementia with Lewy bodies: a phase III, multicentre study. Lancet Neurol 2007;6(4):305–13

19.Thomas AJ, Attems J, Colloby SJ, O'Brien JT, McKeith I, Walker R, Lee L, Burn D, Lett DJ, Walker Z. Autopsy validation of 123I-FP-CIT dopaminergic neuroimaging for the diagnosis of DLB. Neurology. 2017 Jan 17;88(3):276-283

20. McKeith IG, Boeve BF, Dickson DW, Halliday G, Taylor JP, Weintraub D, Aarsland D, Galvin J, Attems J, Ballard CG, Bayston A, Beach TG, Blanc F, Bohnen N, Bonanni L, Bras J, Brundin P, Burn D, Chen-Plotkin A, Duda JE, El-Agnaf O, Feldman H, Ferman TJ, Ffytche D, Fujishiro H, Galasko D, Goldman JG, Gomperts SN, Graff-Radford NR, Honig LS, Iranzo A, Kantarci K, Kaufer D, Kukull W, Lee VMY, Leverenz JB, Lewis S, Lippa C, Lunde A, Masellis M, Masliah E, McLean P, Mollenhauer B, Montine TJ, Moreno E, Mori E, Murray M, O'Brien JT, Orimo S, Postuma RB, Ramaswamy S, Ross OA, Salmon DP, Singleton A, Taylor A, Thomas A, Tiraboschi P, Toledo JB, Trojanowski JQ, Tsuang D, Walker Z, Yamada M, Kosaka K. Diagnosis and management of dementia with Lewy bodies: Fourth consensus report of the DLB Consortium. Neurology. 2017 Jul 4;89(1):88-100

21. Lorberboym M, Treves TA, Melamed E, Lampl Y, Hellmann M, Djaldetti R. [123I]-FP/CIT SPECT imaging for distinguishing drug-induced parkinsonism from Parkinson’s disease. Mov Disord. 2006;21:510–514.

22. Felicio AC, Godeiro-Junior C, Moriyama TS, et al. Degenerative parkinsonism in patients with psychogenic parkinsonism: a dopamine transporter imaging study. Clin Neurol Neurosurg. 2010;112:282–285.

23. Brigo F, Matinella A, Erro R, Tinazzi M. [¹²³I]FP-CIT SPECT (DaTSCAN) may be a useful tool to differentiate between Parkinson's disease and vascular or drug-induced parkinsonisms: a meta-analysis. Eur J Neurol. 2014Nov;21(11):1369-e90.

24. Bauckneht M, Chincarini A, De Carli F, Terzaghi M, Morbelli S, Nobili F, Arnaldi D. Presynaptic dopaminergic neuroimaging in REM sleep behavior disorder: A systematic review and meta-analysis. Sleep Med Rev. 2018 Oct;41:266-274

25. Berg D, Postuma RB, Adler CH, Bloem BR, Chan P, Dubois B, Gasser T, Goetz CG, Halliday G, Joseph L, Lang AE, Liepelt-Scarfone I, Litvan I, Marek K, Obeso J, Oertel W, Olanow CW, Poewe W, Stern M, Deuschl G. MDS research criteria for prodromal Parkinson's disease. Mov Disord. 2015 Oct;30(12):1600-11

26. Schreckenberger M, Hagele S, Siessmeier T, et al. The dopamine D2 rec ept or li gan d 18F -d es met hoxy fa ll ypr ide: an appropriate fluorinated PET tracer for the differential diagnosis of parkinsonism. Eur J Nucl Med Mol Imaging. 2004;31:1128–35.

27. Hellwig S, Amtage F, Kreft A, Buchert R, Winz OH, Vach W, Spehl TS, Rijntjes M, Hellwig B, Weiller C, Winkler C, Weber WA, Tüscher O, Meyer PT. [¹⁸F]FDG-PET is superior to [¹²³I]IBZM-SPECT for the differential diagnosis of parkinsonism. Neurology. 2012 Sep 25;79(13):1314-22

28. Yoshita, M. Differentiation of idiopathic Parkinson’s disease from striatonigral degeneration and progressive supranuclear palsy using iodine-123 meta-iodobenzylguanidine myocardial scintigraphy. J Neurol Sci 1998 155, 60-67.

29. Oertel WH, Tatsch K, Schwarz J, et al. Decrease of D2 receptors indicated by 123I-iodobenzamide single-photon emission computed tomography relates to neurological deficit in treated Wilson’s disease. Ann Neurol. 1992;32:743–8

30.Kaasinen V, Vahlberg T. Striatal dopamine in Parkinson disease: A meta-analysis of imaging studies Ann Neurol. 2017 Dec;82(6):873-882.

31. Dopaminergic function and dopamine transporter binding assessed with positron emission tomography in Parkinson disease. Ribeiro MJ, Vidailhet M, Loc'h C, Dupel C, Nguyen JP, Ponchant M, Dollé F, Peschanski M, Hantraye P, Cesaro P, Samson Y, Remy P. Arch Neurol. 2002 Apr;59(4):580-

32. Eshuis SA, Maguire RP, Leenders KL, Jonkman S, Jager PL. Comparison of FP-CIT SPECT with F-DOPA PET in patients with de novo and advanced Parkinson's disease. Eur J Nucl Med Mol Imaging. 2006 Feb;33(2):2009.

33. Booij J, Kemp P. Dopamine transporter imaging with [(123)I]FPCIT SPECT: potential effects of drugs. Eur J Nucl Med Mol Imaging 2008;35(2):424–38.

34. Volkow ND1, Fowler JS, Logan J, Alexoff D, Zhu W, Telang F, Wang GJ, Jayne M, Hooker JM, Wong C, Hubbard B, Carter P, Warner D, King P, Shea C, Xu Y, Muench L, Apelskog-Torres K Effects of modafinil on dopamine and dopamine transporters in the male human brain: clinical implications. JAMA. 2009 Mar 18;301(11):1148-54.

35. Iffland K, Grotenhermen F. An Update on Safety and Side Effects of Cannabidiol: A Review of Clinical Data and Relevant Animal Studies. CannabisCannabinoid Res. 2017 Jun 1;2(1):139-154. doi: 10.1089/can.2016.0034. eCollection 2017.

36. Booij J, de Jong J, de Bruin K, Knol RJJ, de Win MM, van Eck- Smit BLF. Quantification of striatal dopamine transporters with [123I]FP-CIT SPECT is influenced by the selective serotonin reuptake inhibitor paroxetine: a double-blind, placebo-controlled, crossover study in healthy controls. J Nucl Med 2007;48:359–66

37. Cilia R, Marotta G, Belletti A, Siri C, Pezzoli G. Reversible dopamine transporter reduction in drug-induced Parkinsonism. Mov Disord. 2014 Apr;29(4):575-746.

38. Schillaci O, Pierantozzi M, Filippi L, et al. The effect of levodopa therapy on dopamine transporter SPECT imaging with 123I-FP-CIT in patients with Parkinson’s disease. Eur J Nucl Med Mol Imaging. 2005;32:1452–1456.

39. Yang YK, Yao WJ, Yeh TL, Lee IH, Chen PS, Lu RB, et al. Decreased dopamine transporter availability in male smokers-a dual isotope SPECT study. Prog Neuropsychopharmacol Biol Psychiatry 2008;32(1):274–9.

40. Seibyl JP, Wallace E, Smith EO, Stabin M, Baldwin RM, Zoghbi S, et al. Whole-body biodistribution, radiation absorbed dose and brain SPECT imaging with iodine-123-beta-CIT in healthy human subjects. J Nucl Med 1994;35(5):764–70.

41. Varrone A, Sansone V, Pellecchia MT, Amboni M, Salvatore E, De Michele G, et al. Comparison between a dual-head and a brain-dedicated SPECT system in the measurement of the loss of dopamine transporters with [(123)I]FP-CIT. Eur J Nucl Med Mol Imaging 2008;35(7):1343–

42. Dobbeleir AA, Hambÿe A-SE, Franken PR. Influence of highenergy photons on the spectrum of iodine-123 with low-and medium-energy collimators: consequences for imaging with 123Ilabelled compounds in clinical practice. Eur J Nucl Med Mol Imaging 1999;26(6):655

43. Koch W, Bartenstein P, la Fougère C. Radius dependence of FP-CIT quantification: a Monte Carlo-based simulation study. Ann Nuc Med. 2014;28:103–11.

44. QIBA Profile: Quantifying Dopamine Transporters with 123 - Iodine Labeled Ioflupane in Neurodegenerative Disease Profile (v1.1), August 28, 2017

45. Dickson, J.C. et al., 2009. The impact of reconstruction method on the quantification of DaTSCAN images. European journal of nuclear medicine and molecular imaging, 37(1), pp.23–35.

46. Tossici-Bolt, L. et al., 2011. Calibration of gamma camera systems for a multicentre European 123I-FP-CIT SPECT normal database. European journal of nuclear medicine and molecular imaging, 38(8), pp.1529–1540.

47. Soret, M. et al., 2003. Quantitative accuracy of dopaminergic neurotransmission imaging with (123)I SPECT. *Journal of Nuclear Medicine*, 44(7), pp.1184–1193

48. Ichihara, T. et al., 1993. Compton scatter compensation using the triple-energy window method for single- and dual-isotope SPECT. Journal of Nuclear Medicine, 34(12), pp.2216–2221.

49. Iida, H. et al., 1998. Effects of scatter and attenuation correction on quantitative assessment of regional cerebral blood flow with SPECT. Journal of Nuclear Medicine, 39(1), pp.181–189

50. Tossici-Bolt, L. et al., 2017. [(123)I]FP-CIT ENC-DAT normal database: the impact of the reconstruction and quantification methods. EJNMMI Physics, 4(1), p.8.

51. EANM Physics Committee et al., 2010. Routine quality control recommendations for nuclear medicine instrumentation. European journal of nuclear medicine and molecular imaging, 37(3), pp.662–671.

52. INTERNATIONAL ATOMIC ENERGY AGENCY, Quality Assurance for SPECT Systems, IAEA Human Health Series No. 6, IAEA, Vienna (2009).

53. INTERNATIONAL ATOMIC ENERGY AGENCY, Quality Assurance Programme for Computed Tomography: Diagnostic and Therapy Applications, IAEA Human Health Series No. 19, IAEA, Vienna (2012)

54. Tissingh G, Booij J, Bergmans P, et al. Iodine-123-N-omega-fluoropropyl-2 bcarbomethoxy-3 b-(4-iodophenyl)tropane SPECT in healthy controls and early stage, drug-naive Parkinson’s disease. J Nucl Med. 1998;39:1143–1148.

55. Marshall VL, Reininger CB, Marquardt M, et al. Parkinson’s disease is overdiagnosed clinically at baseline in diagnostically uncertain cases: a 3-year European multicenter study with repeat [123I]FP-CIT SPECT. Mov Disord. 2009;24:500–508.

56. Catafau AM, Tolosa E. DaTSCAN Clinically Uncertain Parkinsonian Syndromes Study Group. Impact of dopamine transporter SPECT using 123I-ioflupane on diagnosis and management of patients with clinically uncertain parkinsonian syndromes. Mov Disord. 2004;19:1175–1182.

57. Arnaldi D, De Carli F, Picco A, Ferrara M, Accardo J, Bossert I, Famà F, Girtler N, Morbelli S, Sambuceti G, Nobili F. Nigro-caudate dopaminergic deafferentation: a marker of REM sleep behavior disorder? Neurobiol Aging. 2015 Dec;36(12):3300-3305

58. Pirker W, Holler I, Gerschlager W, Asenbaum S, Zettinig G, Brücke T. Measuring the rate of progression of Parkinson's disease over a 5-year period with beta-CIT SPECT. Mov Disord. 2003 Nov;18(11):1266-72

59. Booij J, Dubroff J, Pryma D, Yu J, Agarwal R, Lakhani P, Kuo PH. Diagnostic Performance of the Visual Reading of (123)I-Ioflupane SPECT Images With or Without Quantification in Patients With Movement Disorders or Dementia. J Nucl Med. 2017 Nov;58(11):1821-1826

60. Marek K, Seibyl J, Eberly S, Oakes D, Shoulson I, Lang AE, Hyson C, Jennings D; Parkinson Study Group PRECEPT Investigators. Longitudinal follow-up of SWEDD subjects in the PRECEPT Study. Neurology. 2014 May 20;82(20):1791-7

61. Nicastro N, Garibotto V, Badoud S, Burkhard PR. Scan without evidence of dopaminergic deficit: A 10-year retrospective study. Parkinsonism Relat Disord. 2016 Oct;31:53-58.

62. Lorberboym M, Djaldetti R, Melamed E, Sadeh M, Lampl Y. 123I-FP-CIT SPECT imaging of dopamine transporters in patients with cerebrovascular disease and clinical diagnosis of vascular parkinsonism. J Nucl Med. 2004;45:1688–1693.

63. Benítez-Rivero S, Marín-Oyaga VA, García-Solís D, Huertas-Fernández I, García-Gómez FJ, Jesús S, Cáceres MT, Carrillo F, Ortiz AM, Carballo M, Mir P. Clinical features and 123I-FP-CIT SPECT imaging in vascular parkinsonism and Parkinson's disease. J Neurol Neurosurg Psychiatry. 2013 Feb;84(2):122-9

64. Vizcarra JA, Lang AE, Sethi KD, Espay AJ. Vascular Parkinsonism: deconstructing a syndrome. Mov Disord. 2015;30:886-894.

65. Wardlaw JM, Smith EE, Biessels GJ, Cordonnier C, Fazekas F, Frayne R, Lindley RI, O’Brien JT, Barkhof F, Benavente OR, Black SE, Brayne C, Breteler M, Chabriat H, Decarli C, de Leeuw FE, Doubal F, Duering M, Fox NC, Greenberg S, Hachinski V, Kilimann I, Mok V, Oostenbrugge R, Pantoni L, Speck O, Stephan BC, Teipel S, Viswanathan A, Werring D, Chen C, Smith C, van Buchem M, Norrving B, Gorelick PB, Dichgans M. Neuroimaging standards for research into small vessel disease and its contribution to ageing and neurodegeneration. Lancet Neurol. 2013;12:822-838.

66. Jellinger KA. The neuropathologic diagnosis of secondary parkinsonian syndromes. Adv. Neurol. 1996; 69:293–303. [PubMed: 8615141]

67. Van der Zande, J.; Booij, J.; Scheltens, P.; Raijmakers, P.; Lemstra, A.W. [123]FP-CIT SPECT scans initially rated as normal became abnormal over time in patients with probable dementia with Lewy bodies. Eur J Nucl Med Mol Imaging 2016 43(6), 1060-1066.

68. Nobili, F.; Arnaldi, D.; Morbelli, S. Is dopamine transporter invariably impaired at the time of diagnosis in dementia with Lewy bodies? Eur J Nucl Med Mol Imaging 2016 43(6), 1056-1059.

69. Cilia R, Rossi C, Frosini D, Volterrani D, Siri C, Pagni C, Benti R, Pezzoli G, Bonuccelli U, Antonini A, Ceravolo R. Dopamine Transporter SPECT Imaging in Corticobasal Syndrome. PLoS One. 2011 May 2;6(5):e18301

70. Walker Z, Gandolfo F, Orini S, Garibotto V, Agosta F, Arbizu J, Bouwman F, Drzezga A, Nestor P, Boccardi M, Altomare D, Festari C, Nobili F; EANM-EAN Task Force for the recommendation of FDG PET for Dementing Neurodegenerative Disorders. Clinical utility of FDG PET in Parkinson's disease and atypical parkinsonism associated with dementia. Eur J Nucl Med Mol Imaging. 2018 Jul;45(9):1534-1545. doi: 10.1007/s00259-018-4031-2.PMID: 29779045

71. Orimo S, Suzuki M, Inaba A, Mizusawa H.123I-MIBG myocardial scintigraphy for differentiating Parkinson's disease from other neurodegenerative parkinsonism: a systematic review and meta-analysis. Parkinsonism Relat Disord. 2012 Jun;18(5):494-500. doi: 10.1016/j.parkreldis.2012.01.009.PMID: 22321865

72. Rinne JO, Laine M, Kaasinen V, Norvasuo-Heilä MK, Någren K, Helenius H. Striatal dopamine transporter and extrapyramidal symptoms in frontotemporal dementia. Neurology. 2002;58:1489-1493.

73. Morgan S, Kemp P, Booij J et al. Differentiation of frontotemporal dementia from dementia with Lewy bodies using FP-CIT SPECT. J Neurol Neurosurg Psychiatry. 2012;83:1063-1070.

74. Miyoshi M, Shinotoh H, Wszolek ZK et al. In vivo detection of neuropathologic changes in presymptomatic MAPT mutation carriers: a PET and MRI study. Parkinsonism Relat Disord. 2010;16:404-408.

75. Kishore A, Wszolek ZK, Snow BJ et al. Presynaptic nigrostriatal function in genetically tested asymptomatic relatives from the pallido-ponto-nigral degeneration family. Neurology. 1996;47:1588-1590.

76. Tiraboschi, P.; Corso, A.; Guerra, U.P.; Nobili, F.; Piccardo, A.; Calcagni, M.L.; Volterrani, D.; Cecchin, D.; Tettamanti, M.; Antelmi, L.; Vidale, S.; Sacco, L.; Merello, M.; Stefanini, S.; Micheli, A.; Vai, P.; Capitanio, S.; Gabanelli, S.V.; Riva, R.; Pinto, P.; Biffi, A.M.; Muscio, C.; SCILLA Working Group. 123 I-FP-CIT SPECT and 123 I-MIBG myocardial scintigraphy in differentiating dementia with Lewy bodies from other dementias: a comparative study. Ann Neurol 2016 80(3), 368-378.

77. Arnaldi D, Campus C, Ferrara M, Famà F, Picco A, De Carli F, Accardo J, Brugnolo A, Sambuceti G, Morbelli S, Nobili F. What predicts cognitive decline in de novo Parkinson's disease? Neurobiol Aging. 2012 Jun;33(6):1127.e11-20.

78. Seibyl JP, Marek K, Sheff K, et al. Iodine-123-beta-CIT and iodine-123-FPCIT SPECT measurement of dopamine transporters in healthy subjects and Parkinson’s patients. J Nucl Med. 1998;39:1500–1508.

79. Laruelle, M., et al., Graphical, kinetic, and equilibrium analyses of in vivo [123I] beta-CIT binding to dopamine transporters in healthy human subjects. J Cereb Blood Flow Metab, 1994. 14(6): p. 982-94.

80. Seibyl, J.P., et al., Test/retest reproducibility of iodine-123-betaCIT SPECT brain measurement of dopamine transporters in Parkinson's patients. J Nucl Med, 1997. 38(9): p. 1453-9.

81. Booij, J, Dubroff J, Pryma D, Yu JQ, Agarwal R, Lakhani P and Kuo PH: Diagnostic performance of the visual reading of (123)I-ioflupane SPECT images when assessed with or without quantification in patients with movement disorders or dementia. Journal of Nuclear Medicine. 2017 Nov; 58(11): 1821-1826. http://doi.org/10.2967/jnumed.116.189266

82. Booij J, Habraken JB, Bergmans P, et al. Imaging of dopamine transporters with iodine-123-FP-CIT SPECT in healthy controls and patients with Parkinson’s disease. J Nucl Med. 1998;39:1879–1884.

83. Calvini P, Rodriguez G, Inguglia F, Mignone A, Guerra UP, Nobili F. The basal ganglia matching tools package for striatal uptake semi-quantification: description and validation. Eur J Nucl Med Mol Imaging. 2007 Aug;34(8):1240-53

84. Matesan M, Gaddikeri S, Longfellow K, Miyaoka R, Elojeimy S, Elman S, Hu SC, Minoshima S, Lewis D. I-123 DaTscan SPECT Brain Imaging in Parkinsonian Syndromes: Utility of the Putamen-to-Caudate Ratio. J Neuroimaging. 2018 Nov;28(6):629-634

85. Zubal IG, Early M, Yuan O, Jennings D, Marek K, Seibyl JP. Optimized, automated striatal uptake analysis applied to SPECT brain scans of Parkinson's disease patients. J Nucl Med. 2007 Jun;48(6):857-64

86. Morton RJ, Guy MJ, Clauss R, Hinton PJ, Marshall CA, Clarke EA. Comparison of different methods of DatSCAN quantification. Nucl Med Commun. 2005 Dec;26(12):1139-46

87. Nobili, F., et al., Automatic semi-quantification of [123I]FP-CIT SPECT scans in healthy volunteers using BasGan version 2: results from the ENC-DAT database. Eur J Nucl Med Mol Imaging, 2013. 40(4): p. 565-7

88: Kuo PH, Lei HH, Avery R, Krupinski EA, Bauer A, Sherman S, McMillan N, Seibyl J, Zubal GI: Evaluation of an Objective Striatal Analysis Program for Determining Laterality in Uptake of ¹²³I-Ioflupane SPECT Images: Comparison to Clinical Symptoms and to Visual Reads. J Nucl Med Technol. 2014 Jun;42(2):105-8. doi: 10.2967/jnmt.113.134940.

89. Dickson JC, Tossici-Bolt L, Sera T, et al. The impact of reconstruction method on the quantification of DaTSCAN images. Eur J Nucl Med Mol Imaging. 2010; 37:23–35.

90. Koch W, Radau PE, Munzing W, Tatsch K. Cross-camera comparison of SPECT measurements of a 3-D anthropomorphic basal ganglia phantom. Eur J Nucl Med Mol Imaging. 2006;33:495–502.

91. Semiquantitative Analysis of Dopamine Transporter Scans in Patients With Parkinson Disease. Tinaz S, et. al. Clin Nucl Med. 2018 Jan;43(1).

92. Kuo PH, Eshghi N, Tinaz S, Blumenfeld H, Louis ED, Zubal G. Optimization of Parameters for Quantitative Analysis of 123I-ioflupane SPECT Images for Monitoring of Progression of Parkinson's Disease. J Nucl Med Technol. August 2018:jnmt.118.213181. doi:10.2967/jnmt.118.213181.

93. Tossici-Bolt L, Hoffmann SM, Kemp PM, Mehta RL, Fleming JS. Quantification of [(123)I]FP-CIT SPECT brain images: an accurate technique for measurement of the specific binding ratio. Eur J Nucl Med Mol Imaging. 2006;33:1491–9

94. Hsu CC, Chang YH, Lin WC, Tang SW, Wang PW, Huang YC, Chiu NT. The feasibility of using CT-guided ROI for semiquantifying striatal dopamine transporter availability in a hybrid SPECT/CT system. ScientificWorldJournal. 2014;2014:879497

95. Hirschauer, T.J., H. Adeli, and J.A. Buford, Computer-Aided Diagnosis of Parkinson's Disease Using Enhanced Probabilistic Neural Network. J Med Syst, 2015. 39(11): p. 179.

96. Rahmim, A., et al., Application of texture analysis to DAT SPECT imaging: Relationship to clinical assessments. Neuroimage Clin, 2016. 12: p. e1-e9.

97. Mozley PD, Schneider JS, Acton PD, Plössl K, Stern MB, Siderowf A, Leopold NA, Li PY, Alavi A, Kung HF. Binding of [99mTc]TRODAT-1 to dopamine transporters in patients with Parkinson's disease and in healthy volunteers. J Nucl Med. 2000 Apr;41(4):584-9..

98. Dickson JC, Tossici-Bolt L, Sera T, Booij J, Ziebell M, Morbelli S, Assenbaum-Nan S, Borght TV, Pagani M, Kapucu OL, Hesse S, Van Laere K, Darcourt J, Varrone A, Tatsch K. The impact of reconstruction and scanner characterization on the diagnostic capability of a normal database for [(123)I]FP-CIT SPECT imaging. EJNMMI Res. 2017 Dec;7(1):1

99. Buchert R, Kluge A, Tossici-Bolt L, Dickson J, Bronzel M, Lange C, Asenbaum S, Booij J, Atay Kapucu LÖ, Svarer C, Koulibaly PM, Nobili F, Pagani M, Sabri O, Sera T, Tatsch K, Vander Borght T, Van Laere K, Varrone A, Iida H. Reduction in camera-specific variability in [(123)I]FP-CIT SPECT outcome measures by image reconstruction optimized for multisite settings: impact on age-dependence of the specific binding ratio in the ENC-DAT database of healthy controls. Eur J Nucl Med Mol Imaging. 2016 Jul;43(7):1323-36

100. Varrone A, Dickson JC, Tossici-Bolt L, Sera T, Asenbaum S, Booij J, Kapucu OL, Kluge A, Knudsen GM, Koulibaly PM, Nobili F, Pagani M, Sabri O, Vander Borght T, Van Laere K, Tatsch K. European multicentre database of healthy controls for [123I]FP-CIT SPECT (ENC-DAT): age-related effects, gender differences and evaluation of different methods of analysis. Eur J Nucl Med Mol Imaging. 2013 Jan;40(2):213-27

101. Parkinson Progression Marker Initiative. The Parkinson Progression Marker Initiative (PPMI). Prog Neurobiol. 2011 Dec;95(4):629-35

102. Hesse S, van de Giessen E, Zientek F, Petroff D, Winter K, Dickson JC, Tossici-Bolt L, Sera T, Asenbaum S, Darcourt J, Akdemir UO, Knudsen GM, Nobili F, Pagani M, Vander Borght T, Van Laere K, Varrone A, Tatsch K, Sabri O, Booij J. Association of central serotonin transporter availability and body mass index in healthy Europeans. Eur Neuropsychopharmacol. 2014 Aug;24(8):1240-7

103. Lapa C, Spehl TS, Brumberg J, Isaias IU, Schlögl S, Lassmann M, Herrmann K, Meyer PT. Influence of CT-based attenuation correction on dopamine transporter SPECT with [(123)I]FP-CIT. Am J Nucl Med Mol Imaging. 2015 Feb 15;5(3):278-86

104. Simuni T, Siderowf A, Lasch S, Coffey CS, Caspell-Garcia C, Jennings D, Tanner CM, Trojanowski JQ, Shaw LM, Seibyl J, Schuff N, Singleton A, Kieburtz K, Toga AW, Mollenhauer B, Galasko D, Chahine LM, Weintraub D, Foroud T, Tosun D, Poston K, Arnedo V, Frasier M, Sherer T, Chowdhury S, Marek K; Parkinson's Progression Marker Initiative\*. Longitudinal Change of Clinical and Biological Measures in Early Parkinson's Disease: Parkinson's Progression Markers Initiative Cohort. Mov Disord. 2018 May;33(5):771-782

105. Dickson, J.C. et al., 2011. Proposal for the standardisation of multi-centre trials in nuclear medicine imaging: prerequisites for a European 123I-FP-CIT SPECT database. European journal of nuclear medicine and molecular imaging, 39(1), pp.188–197.

106. Gründer G, Vernaleken I, Müller MJ, Davids E, Heydari N, Buchholz HG, Bartenstein P, Munk OL, Stoeter P, Wong DF, Gjedde A, Cumming P. Subchronic haloperidol downregulates dopamine synthesis capacity in the brain of schizophrenic patients in vivo. Neuropsychopharmacology. 2003 Apr;28(4):787-94

107. Stout DB, Huang SC, Melega WP, Raleigh MJ, Phelps ME, Barrio JR (1998). Effects of large neutral amino acid concentrations on 6-[F-18]Fluoro-L-DOPA kinetics. J Cereb Blood Flow Metab 18:43–51.

108. Bloomfield MA1, Pepper F1, Egerton A1, Demjaha A1, Tomasi G2, Mouchlianitis E3, Maximen L4, Veronese M5, Turkheimer F5, Selvaraj S3, Howes OD1. Dopamine function in cigarette smokers: an [¹⁸F]-DOPA PET study. Neuropsychopharmacology. 2014 Sep;39(10):2397-404. doi: 10.1038/npp.2014.87. Epub 2014 Apr 10.

109. Cumming, P., Ase, A., Diksic, M., Harrison, J., Jolly, D., Kuwabara, H., Laliberte, C. & Gjedde, A. (1995) Metabolism and blood-brain clearance of L-3,4-dihydroxy-[3H]phenylalanine ([3H]DOPA) and 6-[18F]fluoro-L-DOPA in the rat. Biochem Pharmacol, 50, 943-946

110. Brown WD, Oakes TR, DeJesus OT, Taylor MD, Roberts AD, Nickles RJ, Holden JE. Fluorine-18-fluoro-L-DOPA dosimetry with carbidopa pretreatment. J Nucl Med. 1998 Nov;39(11):1884-91

111. Sawle, G.V., Burn, D.J., Morrish, P.K., Lammertsma, A.A., Snow, B.J., Luthra, S., Osman, S. & Brooks, D.J. (1994) The effect of entacapone (OR-611) on brain [18F]-6-L-fluorodopa metabolism: implications for levodopa therapy of Parkinson's disease. Neurology, 44, 1292-129

112 Dhawan V, Ma Y, Pillai V, Spetsieris P, Chaly T, Belakhlef A, et al. Comparative analysis of striatal FDOPA uptake in Parkinson's disease: ratio method versus graphical approach. J Nucl Med. 2002 Oct;43(10):1324-30.

113. Darcourt J, Schiazza A, Sapin N, Dufour M, Ouvrier MJ, Benisvy D, Fontana X, Koulibaly PM. 18F-FDOPA PET for the diagnosis of parkinsonian syndromes. Q J Nucl Med Mol Imaging. 2014 Dec;58(4):355-65.

114. Jokinen P1, Helenius H, Rauhala E, Brück A, Eskola O, Rinne JO. Simple ratio analysis of 18F-fluorodopa uptake in striatal subregions separates patients with early Parkinson disease from healthy controls. J Nucl Med. 2009 Jun;50(6):893-9.

115. Darcourt J, Schiazza A, Sapin N, Dufour M, Ouvrier MJ, Benisvy D, Fontana X, Koulibaly PM. 18F-FDOPA PET for the diagnosis of parkinsonian syndromes. Q J Nucl Med Mol Imaging. 2014 Dec;58(4):355-65.

116. Whone AL, Watts RL, Stoessl AJ, Davis M, Reske S, Nahmias C, et al. Slower progression of Parkinson's disease with ropinirole versus levodopa: The REAL-PET study. Ann Neurol. 2003 Jul;54(1):93-101.

117. Nagaki A, Onoguchi M, Matsutomo N. Clinical validation of high-resolution image reconstruction algorithms in brain 18F-FDG-PET: effect of incorporating Gaussian filter, point spread function, and time-of-flight. Nucl Med Commun. 2014 Dec;35(12):1224-32.

118. Prieto E, Martí-Climent JM, Morán V, Sancho L, Barbés B, Arbizu J, Richter JA. Brain PET imaging optimization with time of flight and point spread function modelling. Phys Med. 2015 Dec;31(8):948-955.

119. Erlandsson K, Buvat I, Pretorius PH, Thomas BA, Hutton BF. A review of partial volume correction techniques for emission tomography and their applications in neurology, cardiology and oncology. Phys Med Biol. 2012 Nov 7;57(21):R119-59.

120. Rahmim A(1), Qi J, Sossi V. Resolution modeling in PET imaging: theory, practice, benefits, and pitfalls. Med Phys. 2013 Jun;40(6):064301.

121. Sawle GV, Playford ED, Burn DJ, Cunningham VJ, Henderson BT, Brooks DJ (1994) Separating Parkinson’s disease from normality: Discriminant function analysis of [18F]Dopa PET data. Arch.Neurol. 51, 237-243.

122. Takikawa S, Dhawan V, Chaly T, Robeson W, Dahl R, Zanzi I, Mandel F, Spetsieris P, Eidelberg D (1994) Input functions for 6-[fluorine-18]fluorodopa quantitation in parkinsonism: comparative studies and clinical correlations. J Nucl Med 35, 955-963.

123. Pavese N, Rivero-Bosch M, Lewis SJ, Whone AL, Brooks DJ. Progression of monoaminergic dysfunction in Parkinson's disease: a longitudinal 18F-dopa PET study. NeuroImage. 2011 Jun 1;56(3):1463-8 ).

125. Patlak CS, Blasberg RG, Fenstermacher JD (1983) Graphical evaluation of blood-to-brain transfer constants from multiple-time uptake data. J.Cereb.Blood.Flow.Metab. 3, 1-7.

126. Brown, R. M., Crane, A. M., and Goldman, P. S. (1979). Regional distribution of monoamines in the cerebral cortex and subcortical structures of the rhesus monkey: concentrations and in vivo synthesis rates. Brain Res. 168, 133–150. doi: 10.1016/0006-8993(79)90132-x

127. Brooks DJ, Ibañez V, Sawle GV, Quinn N, Lees AJ, Mathias CJ, Bannister R, Marsden CD, Frackowiak RSJ (1990) Differing patterns of striatal 18F-dopa uptake in Parkinson's disease, multiple system atrophy and progressive supranuclear palsy. Ann.Neurol. 28, 547-555.

128. Sossi V, de La Fuente-Fernandez R, Holden JE, Doudet DJ, McKenzie J, Stoessl AJ, Ruth TJ (2002 Feb) Increase in dopamine turnover occurs early in Parkinson's disease: evidence from a new modeling approach to PET 18 F-fluorodopa data. J Cereb Blood Flow Metab 22, 232-239.

129. Kumakura et al. Elevated [18F]fluorodopamine turnover in brain of patients with schizophrenia: an [18F]fluorodopa/positron emission tomography study. 2007 J Neurosci 27:8080-7

130. Boellaard R, O’Doherty MJ, Weber WA, Mottaghy FM, Lonsdale MN, Stroobants SG, et al. FDG PET and PET/CT: EANM procedure guidelines for tumour PET imaging: version 1.0. Eur J Nucl Med Mol Imaging. 2010;37:181–200.

131. Varrone A, Halldin C. New developments of dopaminergic imaging in Parkinson’s disease. Q. J. Nucl. Med. Mol. Imaging. 2012. PMID: 22460161

132. Halldin C, Erixon-Lindroth N, Pauli S, Chou YH, Okubo Y, Karlsson P, et al. [11C]PE21: A highly selective radioligand for PET examination of the dopamine transporter in monkey and human brain. Eur J Nucl Med Mol Imaging. 2003; PMID: 12811422.

133. Jucaite A, Odano I, Olsson H, Pauli S, Halldin C, Farde L. Quantitative analyses of regional [11C]PE2I binding to the dopamine transporter in the human brain: A PET study. Eur J Nucl Med Mol Imaging. 2006; PMID: 16514530

134. Hirvonen J, Johansson J, Teräs M, Oikonen V, Lumme V, Virsu P, et al. Measurement of striatal and extrastriatal dopamine transporter binding with high-resolution PET and [11C]PE2I: Quantitative modeling and test-retest reproducibility. J Cereb Blood Flow Metab. 2008; PMID: 18183028

135. Paul G, Zachrisson O, Varrone A, Almqvist P, Jerling M, Lind G, et al. Safety and tolerability of intracerebroventricular PDGF-BB in Parkinson’s disease patients. J Clin Invest. 2015; PMID: 25689258.

136. Appel L, Jonasson M, Danfors T, Nyholm D, Askmark H, Lubberink M, et al. Use of 11C-PE2I PET in Differential Diagnosis of Parkinsonian Disorders. J Nucl Med. 2015; PMID: 25593112.

137. Li W, Lao-Kaim NP, Roussakis AA, Martín-Bastida A, Valle-Guzman N, Paul G, et al. 11C-PE2I and18F-Dopa PET for assessing progression rate in Parkinson’s: A longitudinal study. Mov Disord. 2018; PMID: 29082547.

138. Inaji M, Okauchi T, Ando K, Maeda J, Nagai Y, Yoshizaki T, et al. Correlation between quantitative imaging and behavior in unilaterally 6-OHDA-lesioned rats. Brain Res. 2005; PMID: 16298352

139. Aznavour N, Cendres-Bozzi C, Lemoine L, Buda C, Sastre JP, Mincheva Z, et al. MPTP Animal Model of Parkinsonism: Dopamine Cell Death or Only Tyrosine Hydroxylase Impairment? - A Study Using PET Imaging, Autoradiography, and Immunohistochemistry in the Cat. CNS Neurosci Ther. 2012; 23106974.

140. DeLorenzo C, Kumar JD, Zanderigo F, Mann JJ, Parsey R V. Modeling considerations for in vivo quantification of the dopamine transporter using 11 CPE2I and positron emission tomography. J Cereb Blood Flow Metab. 2009; PMID: 19458606.

141. Varrone A, Steiger C, Schou M, Takano A, Finnema SJ, Guilloteau D, et al. In vitro autoradiography and in vivo evaluation in cynomolgus monkey of [18F]FE-PE2I, a new dopamine transporter PET radioligand. Synapse. 2009;63:871-80.

142. Varrone A, Toth M, Steiger C, Takano A, Guilloteau D, Ichise M, et al. Kinetic analysis and quantification of the dopamine transporter in the nonhuman primate brain with 11C-PE2I and 18F-FE-PE2I. J Nucl Med. 2011;52:132-9.

143. Sasaki T, Ito H, Kimura Y, Arakawa R, Takano H, Seki C, et al. Quantification of dopamine transporter in human brain using PET with 18F-FE-PE2I. J Nucl Med. 2012;53:1065-73.

144. Fazio P, Svenningsson P, Forsberg A, Jonsson EG, Amini N, Nakao R, et al. Quantitative Analysis of (1)(8)F-(E)-N-(3-Iodoprop-2-Enyl)-2beta-Carbofluoroethoxy-3beta-(4'-Methyl-Phenyl ) Nortropane Binding to the Dopamine Transporter in Parkinson Disease. J Nucl Med. 2015;56:714-20.

145. Sonni I, Fazio P, Schain M, Halldin C, Svenningsson P, Farde L, et al. Optimal Acquisition Time Window and Simplified Quantification of Dopamine Transporter Availability Using 18F-FE-PE2I in Healthy Controls and Parkinson Disease Patients. J Nucl Med. 2016;57:1529-34.

146. Fazio P, Svenningsson P, Cselenyi Z, Halldin C, Farde L, Varrone A. Nigrostriatal dopamine transporter availability in early Parkinson's disease. Mov Disord. 2018;33:592-9.

147. Varrone A, Gulyas B, Takano A, Stabin MG, Jonsson C, Halldin C. Simplified quantification and whole-body distribution of [18F]FE-PE2I in nonhuman primates: prediction for human studies. Nucl Med Biol. 2012;39:295-303.

148. Lizana H, Johansson L, Axelsson J, Larsson A, Ogren M, Linder J, et al. Whole-Body Biodistribution and Dosimetry of the Dopamine Transporter Radioligand (18)F-FE-PE2I in Human Subjects. J Nucl Med. 2018;59:1275-80

149. Jakobson Mo S, Axelsson J, Jonasson L, Larsson A, Ögren MJ, Ögren M, Varrone A, Eriksson L, Bäckström D, Af Bjerkén S, Linder J, Riklund K.Dopamine transporter imaging with [18F]FE-PE2I PET and [123I]FP-CIT SPECT-a clinical comparison. EJNMMI Res. 2018 Nov 15;8(1):10. PMID: 30443684

150. Erickson JD, Eiden LE. Functional identification and molecular cloning of a human brain vescile monoamine transporter. J Neurochem. 1993;61:2314-2317.

151. Scherman D, Raisman R, Ploska A, Agid Y. [3H]Dihydrotetrabenazine, a new in vitro monoaminergic probe for human brain. J Neurochem. 1988;50:1131-1136

152. Schafer MK, Hartwig NR, Kalmbach N, Klietz M, Anlauf M, Eiden LE, et al. Species-specific vesicular monoamine transporter 2 (VMAT2) expression in mammalian pancreatic beta cells: implications for optimising radioligand-based human beta cell mass (BCM) imaging in animal models. Diabetologia. May 2013;56(5):1047-1056.

153. Vander Borght TM, Sima AA, Kilbourn MR, Desmond TJ, Kuhl DE, Frey KA. [3H]methoxytetrabenazine: a high specific activity ligand for estimating monoaminergic neuronal integrity. Neuroscience. 1995;68:955-962.

154. Frey KA, Koeppe RA, Kilbourn MR, Vander Borght TM, Albin RL, Gilman S, et al. Presynaptic monoaminergic vesicles in Parkinson’s disease and normal aging. Ann Neurol. 1996;40:873-884.

155. Bohnen NI, Albin RL, Koeppe RA, Wernette KA, Kilbourn MR, Minoshima S, et al. Positron emission tomography of monoaminergic vesicular binding in aging and Parkinson disease. J Cereb Blood Flow Metab. Jan 18 2006;26:1198-1212.

156. Lin KJ, Weng YH, Wey SP, Hsiao IT, Lu CS, Skovronsky D, et al. Whole-body biodistribution and radiation dosimetry of 18F-FP-(+)-DTBZ (18F-AV-133): a novel vesicular monoamine transporter 2 imaging agent. J Nucl Med. Sep 2010;51(9):1480-1485.

157. Lin KJ, Lin WY, Hsieh CJ, Weng YH, Wey SP, Lu CS, et al. Optimal scanning time window for 18F-FP-(+)-DTBZ (18F-AV-133) summed uptake measurements. Nucl Med Biol. Nov 2011;38(8):1149-1155.

158. Lin KJ, Weng YH, Hsieh CJ, Lin WY, Wey SP, Kung MP, et al. Brain imaging of vesicular monoamine transporter type 2 in healthy aging subjects by 18F-FP-(+)-DTBZ PET. PLoS One. 2013;8(9):e75952.

159. Halldin C, Gulyas B, Lang er O, Farde L. Brain radioligands –state of the art and new trends. Q J Nucl Med. 2001;45:139–52.

160. Pinborg LH, Videbaek C, Knudsen GM, et al. Dopamine D(2) receptor quantification in extrastriatal brain regions using [(123)I] epidepride with bolus/infusion. Synapse. 2000;36:322–9

161. Treglia G, Cason E, Stefanelli A, Cocciolillo F, Di Giuda D, Fagioli G, Giordano A. MIBG scintigraphy in differential diagnosis of Parkinsonism: a meta-analysis. Clin Auton Res. 2012 Feb;22(1):43-55.

162. EU BSS Directive. Council Directive 2013/59/Euratom of 5 december 2013

163. IAEA Safety Guide. Radiation Protection and Safety in Medical Uses of Ionizing Radiation. Specific Safety Guide SSG-46, IAEA, Vienna 2018

164. Mattsson S, Johansson L, Leide Svegborn S, Liniecki J, Noßke D, Riklund KÅ, Stabin M, Taylor D, Bolch W, Carlsson S, Eckerman K, Giussani A, Söderberg L, Valind S; ICRP. Radiation Dose to Patients from Radiopharmaceuticals: a Compendium of Current Information Related to Frequently Used Substances. Ann ICRP. 2015 Jul;44(2 Suppl):7-321-5

165. Kuikka JT, Bergström KA, Ahonen A, Länsimies E. The dosimetry of iodine-123 labelled 2 beta-carbomethoxy-3 beta-(4-iodophenyl)tropane. Eur J Nucl Med. 1994;21(1):53-6

166. Kuikka JT. Dosimetry of iodine-123-beta-CIT. J Nucl Med. 1995 Apr;36(4):711

167. Mozley PD1, Stubbs JB, Plössl K, Dresel SH, Barraclough ED, Alavi A, Araujo LI, Kung HF. Biodistribution and dosimetry of TRODAT-1: a technetium-99m tropane for imaging dopamine transporters. J Nucl Med. 1998 Dec;39(12):2069-76.

168. Kuikka JT, Baulieu JL, Hiltunen J, Halldin C, Bergström KA, Farde L, Emond P, Chalon S, Yu M, Nikula T, Laitinen T, Karhu J, Tupala E, Hallikainen T, Kolehmainen V, Mauclaire L, Maziere B, Tiihonen J, Guilloteau D. Pharmacokinetics and dosimetry of iodine-123 labelled PE2I in humans, a radioligand for dopamine transporter imaging. Eur J Nucl Med. 1998 May;25(5):531-4

169. Ribeiro MJ, Ricard M, Lièvre MA, Bourgeois S, Emond P, Gervais P, Dollé F, Syrota A. Whole-body distribution and radiation dosimetry of the dopamine transporter radioligand [(11)C]PE2I in healthy volunteers. Nucl Med Biol. 2007 May;34(4):465-70.

170. Murthy R, Harris P, Simpson N, Van Heertum R, Leibel R, Mann JJ, Parsey R. Whole body [11C]-dihydrotetrabenazine imaging of baboons: biodistribution and human radiation dosimetry estimates. Eur J Nucl Med Mol Imaging. 2008 Apr;35(4):790-7

# Approval

……. (to be completed once final approval by the Boards will be obtained).

**Table 1**: Medications and drugs abuse which may significantly influence the visual and quantitative analysis of [123I]FP-CIT SPECT

|  |  |  |
| --- | --- | --- |
| **Drug class** | **Drug name** | **Comments\*** |
| Cocaine |  | May decrease striatal [123I]FP-CIT binding (2 days) |
| Amphetamines | d-Amphetamine,methamphetamine,methylphenidate | May decrease striatal [123I]FP-CIT binding (7 days) |
| CNS stimulants | Phentermine or ephedrines | May decrease striatal [123I]FP-CIT binding; influences are likely when used as tablets (1 day) |
| Modafinil |  | May decrease striatal [123I]FP-CIT binding (3 days) |
| Antidepressants | Mazindol, bupropion, radafaxine | May decrease striatal [123I]FP-CIT binding (3 days for mazindol; 8 day for bupropion) |
| Adrenergic agonists | Phenylephrine or norepinephrine | May increase striatal [123I]FP-CIT binding; influences are likely when infused at high doses |
| Anticholinergic drugs |  | Benztropine may decrease striatal binding ratios; other anticholinergics may increase these ratios which will likely not affect visual assessments |
| Opioids | Fentanyl | May decrease striatal [123I]FP-CIT binding (5 days) |
| Anesthetics | Ketamine, PCP, isoflurane | May decrease striatal [123I]FP-CIT binding; of interest particularly for animal SPECT studies, although ketamine and PCP are sometimes used illicitly (1 day) |

\* In brackets time interval approximately equal to 5 half-lives.

**Table 2: Overview of DAT SPECT\* findings in patients with parkinsonism**

|  |  |  |
| --- | --- | --- |
| **Disease** | **Findings** | **References** |
| Idiopatic Parkinson Disease | A *dot-shape* is typical since the earliest (and pre-motor) stages of disease. The putamen (in particular the posterior part) of the most affected hemisphere is more severely affected than the caudate nucleus;the putamen of the less affected hemisphere tends to be early involved (often before the caudate nucleus of the most affected side) | 18, 55-58 |
| Essential tremorPsychogenic parkinsonismDrug-induced parkinsonismAlzheimer’s Disease | Normal *Comma-shape* | 22, 23, 55 |
| Vascular parkinsonism | Poorly defined clinical condition. In patients with small vascular lesions in the basal ganglia, DAT tracer uptake is normal or only slightly diminished.Definite vascular parkinsonism should include a documented ischemic or hemorrhagic lesions in the substantia nigra or nigrostriatal pathway leading to striatal presynaptic DAT deficiency (see main text). | 62-66 |  |
|  |  |  |
| Dementia with Lewy Body | Reduced uptake substantially overlapping IPD features  | 18-20, 59, 67, 68 |
|  Around 10% of patients with pathologic proven DLB show a normal DAT SPECT at the time of the clinical diagnosis (possibly becoming abnormal after 1.5-2 years) |  |
|  |  |
| Multiple System Atrophy | Reduced uptake, often overlapping IPD features (uptake reduction tend to be more symmetric as compared to IPD) | 59 |
| Progressive Supranuclear Palsy | Reduced uptake, often overlapping IPD features (uptake reduction tend to be more symmetric and to involve the caudate nucleus earlier in disease course as compared to IPD) | 59 |
| Cortico-basal degeneration | Reduced uptake, overlapping IPD features (uptake reduction sometimes more asymmetric as compared to IPD, and often involving both putamen and caudate head) | 69-72 |
| Fronto-temporal Dementia | Reduced uptake in 30-60 % of patients with sporadic FTD (usually less pronounced than in IPD)  | 72-76 |

Note \* Given the greater number of available studies describing normal and abnormal distribution of DAT SPECT with respect to [18F]fluorodopa, the different pathophysiology and timeframes of DAT and DOPA reduction and considering the different resolution of PET and SPECT we have limited the present summary to DAT SPECT data. However, it should be noted that normal distribution and the hereby described abnormal patterns in different parkinsonian syndromes are expected to be very similar for DAT SPECT and [18F]fluorodopa PET (see main text for further details

|  |  |  |
| --- | --- | --- |
|

|  |
| --- |
| **Table 3: Factors affecting Specific Binding Ratio (SBR) in [**123I**]-FP-CIT SPECT Imaging** |

 |  |
| **Biological Factors** |  |
| -Dopamine Transporter Density |  |
| -Age and Gender |  |
| -Pharmacokinetics Factors (rate of uptake, metabolism and elimination of tracers)-Genetic: Allelic Variants of DAT-Drugs competing with the tracer for DAT binding |  |
| **Technical Factors** |  |
|  |  |
| -Patient ability to remain motionless in the camera-Equipment:  Resolution and sensitivity of selected camera  Collimator CT-based vs uniform attenuation map\* |  |
| -Performance drifts overtime- Acquisition Dose extravasations (counts in image\*\*) Time post-injection Head position§ Patient movement Radius of rotation-ReconstructionOsem vs FBPFiltrationAttenuation correction Scatter and septal penetration correction PVE correction-Choice of SBR algorithm§§-Size and Placement (spatial registration) of Regions of Interest (including Background region used) |  |
|  |  |

PVE: Partial Volume Effect; CT: Computed Tomography

\*Relevant in case of ill-defined attenuation map boundaries, significant bed or head holder attenuation or incorrect mu value; \*\*Counts affect SBR at very low count levels (<1 million) as the Regions of Interest (ROI) provide an average signal (unless too small ROIs are used). §Inclusion of the whole striata and, possibly, of the whole brain in the field of view, in accordance with the requirements of the SBR algorithm on striatal and reference regions and spatial registration process §§ Based on direct measurements of counts concentration or on total counts in striatal VOIs

**Table 4: Radiation dosimetry (in adults)**

|  |  |  |  |  |
| --- | --- | --- | --- | --- |
| **Radiotracer** | **Organs receiving the highest absorbed dose** |  **(mGy/MBq)** | **Effective dose (mSv/MBq)** | **Reference** |
| [123I]β-CIT | Lung LiverBasal Ganglia\* | 0.100.090.27\*\* | 0.031-0.04 | 40, 165, 166 |
| [123I]FP-CIT | LiverColonGall bladder wallLung | 0.0850.0590.0440.042 | 0.025 | 164 |
| [99mTc]Trodat-1 | LiverKidney | 0.0470.035 | 0.012 | 167 |
| [18F]-FDOPA  | Urinary bladder wallKidney | 0.300.031 | 0.025 | 164 |
| [123I]-PE2I | Urinary bladder wall | 0.07 | 0.022 | 168 |
| [11C]-PE2I | Urinary bladder wallKidneyStomach wall | 0.0180.0160.014 | 0.0064 | 169 |
| [18F]FE-PE2I | Urinary bladder wallLiver | 0.1190.046 | 0.023 | 148 |
| [11C]DTBZ | Stomach wallKidney | 0.0160.013 | 0.0066  | 170 |
| [18F]FP-DTBZ([18F]AV-133) | PancreasLiverUpper Large Intestine | 0.1530.0720.055 | 0.028 | 156 |

Note: when a review performed by the International Commission for Radiological Protection (ICRP) was available this reference was included (ref 164). In other cases, it should be noted that reference may be related to one single article. It should be noted that the human dosimetry for the tracers presented in the table in most cases is based on studies including a very limited number of patients/volunteers. The (high) precision of the numbers shown does not necessarily reflect the actual accuracy of the information. Uncertainties reported in the literature are usually the variation observed in the material but does not include systematic errors in the methods applied. \*The basal ganglia is normally not considered “an organ” in the ICRP context. The value is given here because it is significantly higher than for other tissues, and posted in the reference cited. \*\* Values in the reference is given separately for the two genders. Here the average is presented.